| About the Measure |  |
| --- | --- |
| Domain | Early Psychosis |
| Measure | Brain Activation and Connectivity |
| Definition | A measure for the standardized administration of functional magnetic resonance imaging (fMRI) during task performance for research and clinical efforts. |
| Purpose | This measure can be used to assess brain activity by detecting changes in blood flow during a task when the brain is in use. |

| About the Protocol |  |
| --- | --- |
| Description of Protocol | This protocol includes recommended technical and analytical procedures for the consistent collection of task functional magnetic resonance imaging (fMRI) data based on commonly employed methods for assessment during cognitive functioning testing. These procedures include identification of key MRI acquisition parameters and critical elements of data preprocessing, as well as key interpretational issues. Specific Instructions: Collect resting-state fMRI first, obtaining 24 minutes of total scanning with good data from 12–18 minutes post-resting state. Link to Resting-state fMRI in PhenX Toolkit: https://www.phenxtoolkit.org/toolkit_content/supplemental_info/ptsd/measures/09%20Resting%20State%20fMRI.doc There are two recommended tasks to perform during scanning: (1) AX-Continuous Performance Test (AX-CPT), from the PhenX measure Context Processing; and (2) Relational and Item-Specific Encoding task (RiSE), from the PhenX measure Relational Encoding and Retrieval. These assessments are the tasks to be performed after collecting the resting-state data, during the fMRI. The Context Processing measure, AX-CPT, is a computerized test, implemented in Eprime by the Cognitive Neuroscience Test Reliability and Clinical applications for Schizophrenia (CNTRACS) consortium, which evaluates goal maintenance in working memory. During AX-CPT, the subject views a series of cues and probe sequences (one at a time) and performs a button press, indicating if the stimulus is the target or non-target. The four types of cue-probe combinations offered are AX, AY, BX, and BY, with AX being the only target (correct option). A total of 144 cues/probes are presented in four blocks of 36, with limited response windows. The majority of the cues/probes presented are AX, thus creating an expectancy for the subject to prepare a target response following a cue that is an A and make a target response when a probe is an X. Link to AX-CPT in PhenX Toolkit: TBD The Relational Encoding and Retrieval measure, RiSE, is a computerized test, developed in Eprime, that evaluates item-specific encoding and relational encoding. The subject views visual objects (pictures) and responds yes/no to whether the objects are “living” (item-specific encoding). The subject then views pairs of visual objects and responds yes/no to whether or not one item can fit into the other (relational encoding). Then the subject views visual objects and indicates whether each object was “old” or “new,” responding by way of confidence level high, medium, or low (item recognition). Lastly, the subject views pairs of visual objects and responds yes/no to whether or not the pairs were presented together in the relational encoding test (associative recognition). Link to RiSE in PhenX Toolkit: TBD |
| Protocol Text | Sequence for task functional magnetic resonance imaging (Task fMRI) TA: 6:1 PAT: Off Voxel size: 3.0 × 3.0 × 3.0 mm Rel. SNR: 1.00 Properties Prio Recon ………………………………………………………........... Off Before measurement After measurement Load to viewer …………………………...…………………………...... On Inline movie …………...…………...……………………………........... Off Auto store images ………………………………………………........... On Load to stamp segments ………………………………………........... On Load images to graphic segments ………...…………………………. Off Auto open inline display ………...……………...……………………... On Start measurement without further preparation ………...…………... Off Wait for user to start ………...………...….……………………........... On Start measurements ………...…………………...…………........... single Routine Slice group 1 Slices ………………………….…………...…………………………. 47 Dist. Factor ……………………….………………………………… 0% Position …..…….……………………………………… R3.0 A3.0 H0.0 Orientation ….……….……………………………………… T > C-12.5 Phase enc. dir. ………………………….……………………… A >> P Rotation ………………..……………………………………… 0.00 deg Phase oversampling ………………………...……………...………… 0% FoV read ………………………...…………...…………………… 216 mm FoV phase ……………………………………...………………… 100.0% Slice thickness ……………….……..…………………………… 3.00 mm TR ………………….……………………………………………… 3000 ms TE …………...….……………………….…………………………… 30 ms Averages ……………….……………….………………………………… 1 Concatenations ………………………….……………………..………… 1 Filter ………………………….…………………….… Prescan Normalize Coil elements …………………………...……………………… HEA;HEP Contrast MTC ………………………………...…………………………………… Off Flip angle …………...……….……...……………………………… 85 deg Fat suppr. ……….………….……………………………………… Fat sat. - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - Averaging mode …………………..…………………………… Long term Reconstruction ………………………………………………… Magnitude Measurements ………………………….………………...…………… 124 Delay in TR ………………………….…...…………………...……… 0 ms Multiple series ………………………….…………..………………...… Off Resolution Base resolution ……………………..…………………...……………… 72 Phase resolution ………………………….……………………...… 100% Phase partial Fourier ………………………….……..………………… Off Interpolation ………………………….……….………………………… Off - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - PAT mode …………………….……………………………………… None Matrix Coil Mode ……………….…………………….………… Auto (CP) - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - Distortion Corr. ………………….……………………………………… Off Unfiltered images ………………………….…………………………… Off Prescan Normalize ………………………….………………….……… On Raw filter ………………………….……………………………..……… On Elliptical filter ………………………….………………………………… Off Hamming ………………………….……………………………..……… Off Geometry Multi-slice mode ………………….…………………………… Interleaved Series ………………......……………………………………… Interleaved - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - Special sat. ………………………….……………………..………… None System Body ………………………….……………………………..…………… Off HEP ………………………….……………………………..…………… On HEA ………………………….……………………………..…………… On - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - Positioning mode ………………………….……………………….… REF Table position ………………………….………………………………… H Table position ………………………….……………………….…… 0 mm MSMA ………………………….…………………….…………… S - C - T Sagittal ………………………….…………………...……………… R >> L Coronal ………………………….…………………..……………… A >> P Transversal ………………………….………………...…………… F >> H Coil Combine Mode ……………………….…………… Sum of Squares AutoAlign ………...……………...…………………… Head > Brain Atlas Auto Coil Select ……………………………………...………………… Off - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - Shim mode ………………………….…………………………… Standard Adjust with body coil ………………………….…………….……….… Off Confirm freq. adjustment ………………………….……………...…… Off Assume Silicone ………………………….…………..……………...… Off ? Ref. amplitude 1H ………………………….………...………… 0.000 V Adjustment Tolerance ………………………….………………….… Auto Adjust volume Position ………………………….…………………..… R3.0 A3.0 H0.0 Orientation ……………………..…………………………… T > C-12.5 Rotation …………….….……………………………………… 0.00 deg R >> L …………………..……………………………………… 216 mm A >> P ……………….….……………………………………… 216 mm F >> H ……………………….……….………………………… 141 mm Physio 1st Signal/Mode ………………………….………………………..… None BOLD GLM Statistics ………………………….…….………………………… Off Dynamic t-maps ………………………….…..………………………… Off Starting ignore meas ………………………….………….……………… 0 Ignore after transition ………………………….……………………....… 0 Model transition states ………………………….…...………………… Off Temp. highpass filter ………………………….……………………..… Off Threshold …...………….…………………………………………...… 4.00 Paradigm size ……………….……………………………………….…… 3 Meas[1] ………………………………………………………...… Baseline Meas[2] …………...….…………………………………...……… Baseline Meas[3] ……...……….……………………………………………… Active Motion correction …….………………………………………………… Off Spatial filter ……………...……………………………………………… Off Sequence Introduction ………………..…….……………………………………… Off Bandwidth …….……….……………………………………… 2240 Hz/Px Free echo spacing ……………….………..…………………………… Off Echo spacing ……………………….…………..………………… 0.51 ms - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - EPI factor …………………….………..………………………………… 72 RF pulse type …………………………………...………………… Normal Gradient mode ………………………….…………..………………… Fast - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - Dummy Scans ………………………….………………………………… 0 FFT Scale Factor ………………………….……..………………...… 1.00 |
| Participant | Adolescents and adults, ages 12 and older |
| Rationale for Inclusion in Supplemental Information | Task functional magnetic resonance imaging (task fMRI) provides unique and valuable information about brain activation and connectivity in early psychosis. The Early Psychosis Working Group recommended obtaining this dataset wherever possible. However, task fMRI requires specialized expertise and equipment to obtain in reliable and reproducible fashion. Therefore this measure is more appropriate for inclusion as Supplemental Information instead of in the Toolkit. |
| Source | Ragland, J. D., Ranganath, C., Harms, M. P, Barch, D. M., Gold, J. M., Layher, E., . . . Carter, C. S. (2015). Functional and neuroanatomic specificity of episodic memory dysfunction in schizophrenia: A functional magnetic resonance imaging study of the relational and item-specific encoding task. JAMA Psychiatry, 72(9), 909–916. |
| Language of Source | English |
| Personnel and Training Required | All sites collecting functional magnetic resonance imaging (fMRI) data must have a PhD-level MR physicist on site for implementation, assistance with acquisition (not required to be present at every scan, but required for ongoing quality control [QC] and monitoring) and data analysis. The fMRI technician collecting data should be trained in the administration of functional and structural MRI scans with respect to positioning and instructing the subject (e.g., not to move), running the scanner console, and assessing whether appropriate data were collected. Please see the individual neurocognitive assessment protocols for their specific personnel training requirements. |
| Equipment Needs | Task functional magnetic resonance imaging (task fMRI) requires a 3T or higher field-strength MRI scanner. |
| Protocol Type | Complex instrumentation-based assessment |
| Requirements: | Requirements Category; Required (Yes/No):. Major equipment; Yes. Specialized training; Yes. Specialized requirements for biospecimen collection; No. Average time of greater than 15 minutes in an unaffected individual; Yes. |
| General References | Lopez-Garcia, P., Lesh, T. A., Salo, T., Barch, D. M., MacDonald, A. W., III, Gold, J. M., Ragland, J. D., Strauss, M., Silverstein, S. M., & Carter C. S. (2016). The neural circuitry supporting goal maintenance during cognitive control: A comparison of expectancy AX-CPT and dot probe expectancy paradigms. Cognitive, Affective & Behavioral Neuroscience, 16(1), 164–175. doi:10.3758/s13415-015-0384-1 Poppe, A. B., Barch, D. M., Carter, C. S., Gold, J. M., Ragland, J. D., Silverstein, S. M., & MacDonald, A. W., III. (2016). Reduced frontoparietal activity in schizophrenia is linked to a specific deficit in goal maintenance: A multisite functional imaging study. Schizophrenia Bulletin. pii: sbw036. [Epub ahead of print] |
| Process and Review | N/A – for External Review Panel |