Clinical Center Manual of Procedures

Version 5.0

June 15, 2009

Administrative Core

National Jewish Health

Data Coordinating Center

National Jewish Health

Genotyping and Analysis Core

University of Colorado Health Sciences Center

Brigham and Women’s Hospital/Harvard School of Public Health

Johns Hopkins School of Medicine

Imaging Core

National Jewish Health

University of Iowa

Brigham and Women’s Hospital

Pulmonary Function Core

LDS Hospital

Clinical Centers

Brigham and Women’s Hospital and Fallon Clinic

Baylor College of Medicine and Veterans Administration Medical Center Houston

Columbia University Medical Center

Duke University Medical Center

Johns Hopkins School of Medicine and Bloomberg School of Public Health

Los Angeles Biomedical Research Institute at Harbor-UCLA Medical Center

Morehouse School of Medicine

National Jewish Health

Temple University

University of Alabama, Birmingham

University of California, San Diego

University of Iowa

University of Michigan and Ann Arbor Veterans Administration Medical Center

University of Minnesota and Health Partners Research Foundation

Minnesota Veterans Administration Medical Center

University of Pittsburgh

University of Texas Health Science Center at San Antonio

The Genetic Epidemiology of COPD Program is supported by National

Institutes of Health Grants, U01 HL089897 and U01 HL089856.

COPDGene®

Version 5.0 June 15, 2009

Manual of Operations

–1–

Table of Contents

1.

COPDGene® Overview ......................................................................................................... 5

1a.  Rationale ......................................................................................................................... 5

1b.  Subjects ........................................................................................................................... 5

1c.  Hypotheses and Specific Aims ....................................................................................... 6

1d.  Future Program Goals:.................................................................................................... 7

Human Subject Protection..................................................................................................... 9

2a.  Informed Consent Form.................................................................................................. 9

2b.  Informed Consent Process .............................................................................................. 9

2c.  General Comments on Subject Enrollment..................................................................... 9

Data Coordinating Center.................................................................................................... 12

Subject Enrollment.............................................................................................................. 13

4a.  Before Enrolling Subjects in COPDGene® ................................................................... 13

4b.  Materials Needed for a Study Visit............................................................................... 16

4c.  Subject Identification Number...................................................................................... 17

4d.  Barcode Labels.............................................................................................................. 19

4e.  Informed Consent Administration ................................................................................ 20

4f.  Eligibility Assessment .................................................................................................. 20

Collection of Other Medical Information............................................................................ 28

Subject Enrollment and Evaluation Sequence .................................................................... 29

Questionnaires..................................................................................................................... 31

7a.  Demographic Data and Contact Information................................................................ 32

7b.  Safety Assessment Form............................................................................................... 32

7c.  Medical History Questionnaire ..................................................................................... 33

7d.  Medications Questionnaire ........................................................................................... 33

7e.  Respiratory History Questionnaire ............................................................................... 34

7f.  St. George’s Respiratory Questionnaire ....................................................................... 36

Physical Characteristics....................................................................................................... 37

8a.  Height............................................................................................................................ 37

8b.  Weight........................................................................................................................... 37

8c.  Pulse Oximetry and Heart Rate..................................................................................... 37

Spirometry........................................................................................................................... 38

9a.  Overview....................................................................................................................... 38

9b.  Summary of Measures .................................................................................................. 38

9c.  Setting for Spirometry Testing...................................................................................... 39

9d.  Setting up the ndd EasyOne™ Spirometer ................................................................... 39

9e.  Setting up the Computer to be Used with the Spirometer ............................................ 40

9f.  Medication Use Prior to Testing ................................................................................... 40

9g.  Bronchodilator Administration ..................................................................................... 40

9h.  Contraindications .......................................................................................................... 40

9i.  Conducting the Visit ..................................................................................................... 41

9i.1. Safety checks............................................................................................................... 41

9i.2. Preparing to conduct the spirometry maneuver .......................................................... 41

9i.3. Enter the participant data into the spirometer ............................................................. 42

9i.4. Explain the purpose of the spirometry test.................................................................. 42

Manual of Operations

Table of Contents

2.

3.

4.

5.

6.

7.

8.

9.

COPDGene®

Version5.0 June 15, 2009

–2–

9i.5. Demonstrate the maneuver.......................................................................................... 43

9i.6. Pre-bronchodilator test ................................................................................................ 43

9i.7. Administer the bronchodilator .................................................................................... 43

9i.8. Post-bronchodilator maneuver .................................................................................... 43

9i.9. Print results.................................................................................................................. 43

9i.10. Coaching the participant and troubleshooting problems........................................... 44

9j.  Step by Step Spirometry Instructions ........................................................................... 44

9j.1. Verify EasyOne™ Calibration .................................................................................... 44

9j.2. Conduct the Pre-bronchodilator Spirometry ............................................................... 44

9j.3. Bronchodilator Administration ................................................................................... 45

9j.4. Post-bronchodilator Spirometry .................................................................................. 46

9j.5. Print Your Results ....................................................................................................... 46

9j.6. Upload the Database File to the Website .................................................................... 46

9k.  Acceptable and Reproducible Maneuvers .................................................................... 47

9l.  Reference Values .......................................................................................................... 48

9m. Implementation of Spirometry Quality Control............................................................ 49

9m.1. Factors Affecting Spirometry Quality....................................................................... 49

9n.  Spirometry Calibration, Maintenance, and Hygiene..................................................... 50

9o.  Trouble Shooting .......................................................................................................... 50

9p.  ndd EasyOne™ Configuration Settings for Use in COPDGene® ................................. 50

9q.  Spirometry Checklist .................................................................................................... 51

9r.  Common Questions and Answers for the ndd Spirometer............................................ 52

10.  Six-Minute Walk Test ......................................................................................................... 59

10a.  Contraindications ...................................................................................................... 59

10b.  Facilities/Equipment ................................................................................................. 59

10c.  Walk Course.............................................................................................................. 59

10d.  Procedure .................................................................................................................. 59

10e.  Patient Preparation .................................................................................................... 59

10f. Subject Instructions....................................................................................................... 60

10g.  Instructions during Walking ..................................................................................... 60

10h.  Research Coordinator Responsibilities ..................................................................... 60

10i. Abnormal Test Termination Criteria ............................................................................ 60

11.  Imaging................................................................................................................................ 61

11a.  CT Acquisition at Study Sites................................................................................... 61

11a.1.  Site Survey ................................................................................................................ 61

11a.2.  Technologist Training............................................................................................... 61

11a.3.  Scanner Quality Assurance ....................................................................................... 61

11a.4.  Subject Preparation for CT Scan............................................................................... 62

11a.5.  CT Acquisition.......................................................................................................... 62

11a.6.  Non-study Scans ....................................................................................................... 62

11a.7.  CT Assessment Scoresheet ....................................................................................... 62

11a.8.  Image Acquisition Form ........................................................................................... 63

11a.9.  Scan Shipment .......................................................................................................... 63

11a.10. Quality Assurance..................................................................................................... 65

11b.  Data Reporting .......................................................................................................... 65

11c.  Instructions for CT Scan Acquisition ....................................................................... 66

COPDGene®

Version5.0 June 15, 2009

–3–

Manual of Operations

Table of Contents

12.  Collection of Blood Samples............................................................................................... 69

12a.  Purpose of COPDGene® Study Center LAB Manual ............................................... 69

12b.  Biological Specimen Handling ................................................................................. 69

12b.1.  Labeling and Identification ....................................................................................... 69

12b.2.  Blood Draws/Phlebotomy......................................................................................... 69

12b.3.  Sample Processing .................................................................................................... 70

12c.  Sample Distribution/Shipping................................................................................... 71

12c.1.  Shipper Assembling and Packaging Instructions...................................................... 71

12c.2.  Shipping from COPDGene® Sites to JHBR LAB .................................................... 72

12d.  Supplies and Ordering............................................................................................... 74

12e.  Good Laboratory Practices ....................................................................................... 74

Biorepository Appendices: ........................................................................................................... 76

COPDGene®

Version5.0 June 15, 2009

–4–

Manual of Operations

Table of Contents

1.

1a.

COPDGene® Overview

Rationale

COPD is the fourth leading cause of death in the United States. Of the top ten causes of death,

COPD is the only one that has been steadily increasing over the past decade and thus is one of

the most devastating disease epidemics in the United States. The determinants of the develop-

ment of COPD and the factors that influence its progression remain poorly understood. Although

the most common cause of COPD is cigarette smoking, a minority of people who smoke develop

COPD. Thus, there are individuals who are prone to develop lung disease and other individuals

who are protected from the effects of cigarette smoke. This study will investigate the genetic

determinants associated with the development of COPD.

This study will create a large cohort of subjects who 1) have COPD (GOLD Stages 1 - 4), 2) do

not have COPD but have smoked cigarettes (GOLD 0), 3) Do not have COPD or airway

limitation by FEV1/FVC ratio ≥ 0.70 but present with low FEV1 (<80%) (Unclassified group),

and 4) Do not have COPD and do not smoke with no airflow limitation (non-smoking controls).

This cohort will be phenotyped physiologically (spirometry, 6-minute walk), clinically (health

status, dyspnea, BODE score), and radiographically (chest CT scan). Genome-wide association

analysis will be performed using a staged approach to identify and replicate COPD susceptibility

genes. The cohort will be balanced for gender and age. The size of the cohort has been chosen to

provide sufficient statistical power to identify genetic risk factors in both non-Hispanic White

and non-Hispanic African-American subjects. All results will be validated in the entire cohort

and in two external family-based association cohorts.

The goal of this project is to identify and correlate the genetic risk factors and phenotypic

features that underlie COPD. The cohort will initially be used in a cross-sectional design but is

also designed for future longitudinal follow-up. A series of case-control analyses will be

conducted to identify genes controlling risk for COPD and confirm their effects. The first phase

of this study is a genome-wide screen using a conventional case-control design (stratified by

racial groups). Subsequent confirmation of SNPs (single gene polymorphisms) yielding

statistical signals will be carried out in a second phase of independent case-control analysis.

Finally, we will conduct detailed analysis of genes/regions around these confirmed SNPs that

appear to influence risk of COPD. We will also validate the effects of genes identified here in

two external family-based populations and a large group of mild COPD cases (GOLD Stage 1).

Over time, the phenotypic, pathophysiologic and genetic data collected here can be used to

identify factors that control progression of COPD as these subjects are followed prospectively.

1b.

Subjects

The total number of subjects in COPDGene® is 12,000

1. Subjects with COPD – Up to 6,000 total subjects

• Mild COPD, GOLD Stage 1 – 1,500 subjects

• Moderate to very severe COPD, GOLD Stages 2-4 – Up to 4,500 subjects

COPDGene®

Version5.0 June 15, 2009

–5–

Manual of Operations

Overview

2. Subjects without COPD (history of cigarette smoking and spirometry without airflow

limitation) – Up to 4,500 subjects

3. Subjects without COPD but with airway restriction – Up to 1000 subjects (history of cigarette

smoking and abnormal spirometry FEV1/FVC ratio ≥ 0.70 but present with low FEV1 <80%.

(Please note that these 1000 subjects will be extracted from the COPD smoking Gold 1-4

groups).

4. Subjects with no smoking history and no COPD or airway restriction (normal spirometry and

FEV1 >80%) – Up to 1500 total subjects with 1000 white subjects and 500 African American

subjects.

1c.

Hypotheses and Specific Aims

Hypotheses:

1)Phenotypic characterization of COPD subjects using computed tomography, as well as

clinical and physiological measures, will enable the broad COPD syndrome to be

decomposed into clinically significant subtypes.

2)Genome-wide association studies will identify genetic determinants for COPD

susceptibility that will provide insight into clinically relevant COPD subtypes.

3)Distinct genetic determinants influence the development of emphysema and airway

disease.

To test these hypotheses, we will carry out genome-wide studies for genetic determinants of

COPD that include precise characterization of COPD subjects using computed tomography, as

well as clinical and physiological measures. This high level of detailed phenotypic characteri-

zation will enable the broad COPD syndrome to be decomposed into clinically significant

subtypes and thereby enable more informative searches for genetic associations. To address our

underlying hypotheses, we will complete the following Specific Aims.

Specific Aim 1: Cohort Building

Recruit a large population stratified by severity of COPD, airway restriction, and including

controls (smoking controls - history of cigarette smoking with normal spirometry, non-smoking

controls – no smoking history with normal spirometry) to conduct cross-sectional case-control

studies. We will identify and phenotype COPD cases and control subjects from two racial groups

(non-Hispanic whites and non-Hispanic African Americans) for genetic, epidemiologic, and

natural history studies.

Specific Aim 2: Genome-Wide Association Study

a.Phase 1. A genome-wide panel of single nucleotide polymorphisms (SNPs) will be tested

for associations with COPD in case-control samples of non-Hispanic whites and non-

Hispanic African Americans.

b.Phase 2. Confirmation of SNPs yielding evidence of association in a second case-control

population from these same racial groups to identify genomic regions for intensive

investigation.

COPDGene®

Version5.0 June 15, 2009

Manual of Operations

Overview

–6–

c.

d.

Phase 3. Mapping of 50 genomic regions yielding strong, confirmed association signals

within each racial group in a third case-control population to identify susceptibility genes

for COPD.

Phase 4. Fine mapping of candidate genes to identify susceptibility alleles and/or high risk

haplotypes using

• the entire set of case-control samples from both racial groups collected in Specific

Aim 1.

• external validation using family-based association analysis in the Boston Early-Onset

COPD Study and the International COPD Genetics Network.

Specific Aim 3: Characterization of Subtypes of COPD

a.Characterize unique airway and parenchymal phenotypes among COPD cases and

determine specific clinical, physiologic and functional profiles.

b.Identify susceptibility genes for COPD subtypes, including CT-defined emphysema and

CT-defined airway disease.

Specific Aim 4: Natural history of COPD and Risk Factors for Progression. The cohort will

be established for longitudinal follow-up with regular contact made to determine mortality, co

morbid disease events and disease status based on clinical and/or chest CT evidence of

progression.

1d.

1.

Future Program Goals:

Follow-up this large COPD cohort longitudinally to:

• Determine risk factors that influence which GOLD 1 subjects progress to more severe

stages of COPD

• Assess risk factors for COPD progression, morbidity, and mortality.

Characterize functional variants in the susceptibility genes identified.

Additional external validation in other population cohorts to assess the generalizability of

our findings about genetic determinants of COPD.

2.

3.

Subjects are being enrolled at twenty-one Clinical Centers in the United States:

National Jewish Health, Denver, CO

Brigham and Women’s Hospital, Boston, MA

Baylor College of Medicine, Houston, TX

Columbia University Medical Center, New York, NY

Duke University Medical Center, Durham, NC

Houston Veterans Administration Medical Center, Houston, TX

John Hopkins School of Medicine and Bloomberg School of Public Health, Baltimore, MD

Los Angeles Biomedical Research Institute, Los Angeles, CA

University of Minnesota, Minneapolis, MN

Minnesota Veterans Administration Medical Center, Minneapolis, MN

Health Partners Research Foundation, St. Paul, MN

Morehouse School of Medicine, Atlanta, GA

Temple University School of Medicine, Philadelphia, PA

COPDGene®

Version5.0 June 15, 2009

Manual of Operations

Overview

–7–

University of Alabama, Birmingham, AL

University of California, San Diego, CA

University of Iowa, Iowa City, IA

University of Michigan, Ann Arbor, MI

Ann Arbor Veterans Administration Medical Center, Ann Arbor, MI

University of Pittsburgh, Pittsburgh, PA

University of Texas Health Science Center at San Antonio, TX

Fallon Clinic, Worcester, MA

Please note: The name COPDGene® has been registered. When using this name in printed

material, it should include the registered symbol (®) as a superscript following the name.

COPDGene®

Version5.0 June 15, 2009

–8–

Manual of Operations

Overview

2.

Human Subject Protection

Each site will submit an application to their local governing Institutional Review Board for

Human Subjects (IRB) prior to enrolling any subjects in the study. The most current version of

the study protocol is available on the COPDGene® website

(https://biosweb.njc.org/sec/COPDGene/MainPage.cfm). If required by local IRB regulations, a

copy of the original application to the National Heart, Lung, and Blood Institute may be

submitted along with the IRB application.

A copy of the IRB approval letter and the approved informed consent with the local IRB stamp

of approval must be transmitted to the COPDGene® Administrative Core prior to initiating the

study. Any changes to the consent form or study protocol made and IRB-approved at the Clinical

Center must be sent to the COPDGene® Administrative Core in a tracked changes format as well

as in final clean form. A brief summary of what changes were made should be noted in an email

or accompanying letter. Clinical Centers are not permitted to enroll subjects until receipt of this

information is confirmed by the Administrative Core. Each site should also transmit updated,

stamped informed consent forms and IRB approval letters to the Administrative Core yearly and

whenever the protocol is modified.

2a.

Informed Consent Form

The most current model informed consent form is available on the COPDGene® website

(https://biosweb.njc.org/sec/COPDGene/MainPage.cfm). The informed consent form for each

site may be modified as needed to meet local IRB regulations. Every attempt should be made not

to delete any key concepts and information from the model consent, but rather information may

be added as needed. Questions about IRB issues should be directed to Barry Make, MD, Clinical

Center Director (makeb@njhealth.org; 303-398-1720), at the Administrative Core.

2b.

Informed Consent Process

No procedures can be performed on any subject prior to the subject providing informed consent

and receiving a unique Study Identification Number.

Informed consent is a dynamic interactive process requiring that the subject first fully understand

the protocol and all procedures and only then provide consent for participation. Although the

informed consent process may be delegated to a research coordinator, the local site Principal

Investigator has ultimate responsibility for informed consent.

The subject should be provided ample opportunity to ask questions about the project and

understand the nature of the procedures involved. When possible, a local site investigator should

meet with subjects to answer any questions, review subject understanding of the project, and then

co-sign the consent statement acknowledging subject understanding and consent to participate.

2c.

General Comments on Subject Enrollment

Making people who volunteer for this research project feel good about their experience is

extremely important for the long term success of COPDGene®. We want all subjects who help us

COPDGene®

Version5.0 June 15, 2009

Manual of Operations

Human Subject Protection

–9–

by volunteering to have a sense that their efforts to complete the testing and expend their

valuable time are appreciated by all of us involved in the project.

If subjects feel good about their involvement in the project they will be more likely to encourage

their friends and acquaintances to consider participating. Conversely, if they have a negative

experience they are likely to discourage others from participating. Word of mouth is an excellent

method of recruiting study subjects – especially controls who should be similar to enrolled

COPD subjects in age and socioeconomic status. We suggest that each enrollee identify several

other individuals including spouse and friends who have a history of cigarette smoking. We

would further suggest that enrolled subjects contact such individuals and ask that they contact the

Clinical Center to participate in the study. Informative brochures to assist in recruitment will be

provided by the Administrative Core to each Clinical Center. These recruitment procedures may

require local IRB approval prior to use at the Clinical Centers.

Study coordinators will have the greatest impact on the subjects since they will be contacting the

subjects and spending time with them. There are a number of ways that coordinators and

investigators can maximize the positive aspects of subjects’ encounters with COPDGene®. Many

of these ideas will include your usual behaviors to comply with human subjects protection rules.

The importance of establishing rapport with each subject is of paramount importance to maintain

long-term contact with subjects. This investigation is designed for longitudinal follow-up that

will be the subject of future grant applications. In addition, with subject consent and IRB

approval, these subjects may also be contacted for other future investigations including but not

limited to the NHLBI SPIROMICS initiative.

Following are some techniques that should be used to ensure that subjects have a positive

experience with COPDGene®.

Establish a rapport

• Understand some of the subject’s reasons for participating and address subject’s issues if

appropriate

• Listen empathically to subject’s concerns about symptoms and limitations

• Get some idea of family supports and relationships as you gather information about

contacts

Treat subjects with respect

• Ask how subjects would like to be addressed – Start with formal salutation until told

otherwise

o Keep a record of their preference in your personal notes

• Carry on discussions of health issues in private spaces

• Allow time for questions and explanations

• Recognize or ask about preferences for taking rest breaks, refreshments, scheduling,

company or solitude

• Be aware of cultural preferences and barriers

COPDGene®

Version5.0 June 15, 2009

– 10 –

Manual of Operations

Human Subject Protection

Be honest without setting negative expectations

• Time involved for study participation

• Discomforts and potential side effects

• Necessity to complete all of the parts of the study visit (while recognizing that a study

subject can withdraw at any time)

Be attuned to the unique needs of an older population and the impact of COPD and chronic

illness

• Transportation problems

• Hearing impairments

• Fatigue

• Anxiety

• Shortness of breath

• Effects of age and low oxygen on thinking and processing information

• Mobility limitations

Shape the visit with an eye to the future

• Make it pleasant

• Convey your interest in subjects as individual, unique people and your personal

appreciation of their efforts

• Keep commitments (e.g., to call back or update them about schedule changes)

• Offer all subjects informational COPDGene® IRB-approved brochures about the study to

give to spouse and friends

• Encourage long term participation in COPDGene®

• Explain what the study goals are and the importance of their current and future

participation

• Remind subjects about periodic future contacts

COPDGene®

Version5.0 June 15, 2009

– 11 –

Manual of Operations

Human Subject Protection

3.

Data Coordinating Center

The Data Coordinating Center is located at the National Jewish Medical and Research Center

within the Biostatistics Department headed by James Murphy, PhD. The DCC is under the

direction of Douglas Everett, PhD, Assistant Head of the National Jewish Division of

Biostatistics and Bioinformatics. The DCC develops, maintains and updates the study forms for

data collection, maintains the study website, receives data from the Clinical Centers and all study

Cores, performs quality control of the data, and provides reports to facilitate the conduct of the

study.

Questions about the process of enrollment, use of the website and collection of study-related data

including questionnaires and procedures should be directed to Sandi Uno, DCC Administrative

Assistant at unos@njhealth.org, 303-398-1861. Contact information for other Data Coordinating

Center personnel are on the study website. Questions about the protocol and manual of

operations should be addressed to the COPDGene® Administrative Core Program Manager at

303-270-2399 or Barry Make, MD, Clinical Center Director at makeb@njhealth.org, 303-398-

1720.

Study Website

Essential and supplementary information about COPDGene® is located on the study website. The

home page of the COPDGene® website is:

https://biosweb.njc.org/sec/COPDGene/MainPage.cfm

The COPDGene® study website is used to collect all information, data and tests in the study and

to track transmission of spirometry results, shipment of CT scans, and shipment of blood

samples. Contact information for all study personnel are on the website including Administrative

Core, the Data Coordinating Center, Biorepository Core, Imaging Core, and Pulmonary Function

Core. Contact information for all Clinical Center personnel is also located on the website. All

changes in study personnel including additions and deletions should be reported promptly to the

Administrative Core (email to: COPDGene@njhealth.org) at the National Jewish Medical and

Research Center. Other study information including the current study protocol, model informed

consent, and manual of operations are on the website and updated as necessary.

Each COPDGene® investigator and coordinator is required to have a unique login identity and

password to use the website. Clinical coordinators and investigators should be registered through

the COPDGene® Administrative Core before they request a logon. In order to obtain your logon

identity and password, please contact Ruthie Knowles (KnowlesR@njhealth.org, 303-270-2133)

at the Data Coordinating Center.

COPDGene®

Version5.0 June 15, 2009

– 12 –

Manual of Operations

Data Coordinating Center

4.

4a.

Subject Enrollment

Before Enrolling Subjects in COPDGene®

The following information is provided to assist study personnel who are enrolling subjects in

COPDGene®. It is strongly suggested that these items be carefully considered and addressed

prior to starting the study.

Determine where subjects will be seen for evaluation and testing. Consider the following:

• Privacy for subject

• Quiet location for optimal subject-coordinator communication

• Comfortable seating for subject and coordinator

• Electrical power for computer

• Computer connection to internet

• Proximity to blood drawing

• Proximity to six-minute walk test course

• Proximity to CT scan

• Proximity to oxygen supply if needed for COPD subjects using oxygen

Check to make sure these pieces of equipment are available and located in an area

convenient to where subjects will be seen. Order equipment if needed.

• Stadiometer (preferably measuring height in centimeters, cm)

• Pulse Oximeter

• Blood Pressure Cuff (consider a variety of cuff sizes for larger subjects)

• Scale (preferably measuring weight in kilograms, kg)

Computer (PC): It is highly recommended that you have a computer available during the Study

Visit. (See sections below for information on how to assure inclusion and exclusion criteria are

met and how to complete questionnaires and forms)

• Password-protection is required on computers

• Internet access is necessary to complete on-line “live” Eligibility Assessment and

Assignment forms that assure immediate access to whether subjects meet inclusion and

exclusion criteria

• Internet access is necessary to download and transmit forms, and to download spirometry

results

• A full version of Adobe Acrobat needs to be installed on the study computer. Appropriate

versions include Adobe Acrobat Standard or Professional

• EasyWare™ software for the spirometer should be installed on the computer. It is highly

recommended that spirometry be performed while the spirometer is connected to the

computer. Note: EasyWare™ software for spirometry can only be installed on a PC. All

other study-related procedures and questionnaires can be processed and completed on

either a PC or Mac computer.

• Microsoft Excel™ is the suggested program to locally maintain contact information about

each study subject using a form available on the study website. Alternate forms using other

programs that can capture and store similar information for use only at the Clinical Center

may be used.

COPDGene®

Version5.0 June 15, 2009

Manual of Operations

Subject Enrollment

– 13 –

• A laser printer, or access to a laser printer, is necessary to print barcode labels and study

questionnaires and forms. Laser printers are preferable to ink jet printers

Order Materials and Supplies:

• Paper, pens, clipboard, printer supplies

• NDD EasyOne™ Spirometer Spirettes™ (mouthpieces)

• Nose clips for spirometry

• Albuterol MDI and Aerochamber® spacer

• Stop watch for six-minute walk test

Blood Processing and Shipping Supplies:

• Avery Labels #5267 for printing barcode labels

• Blood drawing supplies: tourniquet, needle, alcohol wipe, band-aid, gloves

• Blood tubes, plastic – 10 ml purple (EDTA) and 10 ml SST

• Centrifuge

• Blood shipping supplies (provided by Johns Hopkins Biorepository Core)

o Blood sample packing boxes and packing materials

o Blood sample shipping boxes

o Cold pack

o Federal Express labels for blood shipment

• Urine Pregnancy Test collection and testing supplies or order forms

Chest CT Scans:

• Have Radiology Department complete and send the CT Scan Site Survey Form to the

Imaging Core prior to scanning any subjects

• Send a CT scan of a custom phantom to the Imaging Core and approval by the Imaging

Core prior to scanning any subjects

• Ensure Radiology Technologists have completed PowerPoint slide training prior to

scanning any subjects. Fax log signed by radiology technologists to Imaging Core to

document completion of training prior to scanning any subjects

• Know the radiologist at your site who will be most involved in COPDGene®

• Know the radiology technologists who will be performing CT scans at your site in

COPDGene®

• Establish a procedure for providing the CT scan forms, DVDs and mailing cases to the

Radiology Department and for retrieving the forms and DVDs following the scan

• Identify a secure location where the back-up CT scan DVD will be stored at your site

• Purchase recommended DVDs

• Obtain Federal Express labels and DVD shipping packages for CT scan shipment from

Imaging Core

COPDGene®

Version5.0 June 15, 2009

– 14 –

Manual of Operations

Subject Enrollment

Spirometry

• Obtain spirometry certification for each coordinator performing spirometry in

COPDGene®

o Perform spirometry on 3 naïve subjects who have not previously performed

spirometry

o Subject IDs for test subjects as part of the spirometry certification process should be

“Cert 1,” “Cert 2,” and “Cert 3”

o Upload certification spirometry to the COPDGene® website

o Obtain confirmation from the Administrative Core that the Pulmonary Function Core

has certified each technician prior to performing spirometry on any COPDGene®

subjects

COPDGene® Study Website

• Become familiar with website use for obtaining study forms, completing Eligibility and

Assessment Forms and downloading subject questionnaires and forms

• Locate contact information for COPDGene® Administrative Core and Data Coordinating

Center personnel and other study personnel

COPDGene® Forms

• Download all current study forms required for a subject visit and store one copy of each on

your computer

• Read and become familiar with all study forms and questionnaires

• Know how to complete all study forms including Eligibility and Assessment Forms

• Know how to transmit study forms and questionnaires to the COPDGene® Data

Coordinating Center

COPDGene® Protocol

• Download the current protocol from website

• Read and know key details of the study

• Understand in detail COPDGene® inclusion and exclusion criteria

Manual of Operations

• Print and read the Manual of Operations and maintain a copy that is readily available

during Study Visits

• Contact the Administrative Core with questions (COPDGene@njhealth.org)

Practice Subject

• Consider using a colleague as a practice subject to become familiar with study procedures

prior to enrolling any subjects in COPDGene®

Locked File Cabinet for Subject Forms and Source Documents

COPDGene®

Version5.0 June 15, 2009

– 15 –

Manual of Operations

Subject Enrollment

4b.

Materials Needed for a Study Visit

Prior to each Study Visit with a COPDGene® subject, you should have ready access to the

following:

• Subject Identification Number – this is a unique number for each study subject; this is

obtained from the COPDGene® website prior to the Study Visit

• Printed barcode labels

• Your local IRB-approved consent form (2 copies)

• Questionnaires and Forms – either printed or on computer except as noted

o Informed Consent and Permissions Form

o Eligibility Form – the Final Eligibility Form must be completed on the computer and

will require an internet connection

o Assessment Form – this form must be completed on the computer and will require an

internet connection

o Safety Assessment Form

o Demographics – Physical Exam Form

o Spirometry Form

o Respiratory Disease History Questionnaire

o Medication History Questionnaire

o Medical History Questionnaire

o St. Georges Respiratory Questionnaire

o SF-36 Form

o 6-Minute Walk Test Form

o Blood Collection Form

o CT Assessment Form with barcode label

o CT Acquisition Form with barcode label

o Subject Contact Information Form (for local use only)

o If previous CT is used for this subject - CT Parameters Form with barcode label

o Discontinuation Form

o Adverse Event Form

• Computer with software: EasyWare™ spirometry software, Adobe Acrobat

• EasyOne™ Spirometer

o Spirettes™

o Nose clip

o Albuterol HFA

o Spacer for albuterol administration, such as Aerochamber®

o Connecting cable from EasyOne™ spirometer to computer

• DVDs (two) for chest CT scan

• Plastic sleeves (two) for chest CT scan DVD

• Stopwatch for walk test

• Oxygen – for subjects using oxygen

• Pens

• Extra paper for notes

• Folder for forms

COPDGene®

Version5.0 June 15, 2009

– 16 –

Manual of Operations

Subject Enrollment

• Blood drawing supplies

o Barcode labels

o Blood tubes, plastic – 10 ml EDTA and 10 ml SST (make sure you have extra tubes

readily available)

o Blood tube adapters

o Alcohol pads

o Tourniquet

o Needles and tubing

o Bandages

o Gauze

o Centrifuge

o Refrigerator for storing samples

o Shipping materials (see Biorepository section for details)

o Blood Collection Form

o Blood Sample Federal Express™ shipping form

• Chest CT scan supplies

o CT order form for your center (varies at each site depending on local procedures)

o Barcode labels

o Permanent marking pen to label DVDs

o 2 DVDs (one is kept at the site as a back-up)

o 2 plastic DVD sleeves

o Chest CT Federal Express™ Shipment Form

4c.

Subject Identification Number

As the required first step, a Subject Identification Number should be obtained. This MUST be

done before a subject is seen for an in-person Study Visit and should be obtained at the time a

preliminary Eligibility Form is completed if used.

Assignment of a Subject Identification Number does not necessarily indicate that the subject has

signed informed consent nor whether the subject has been enrolled in COPDGene®. By requiring

a Subject Identification Number in all potential subjects that are screened, we will be able to

assess the number of subjects that are excluded from enrolling in the study and the reason for

exclusion.

Each subject who participates in screening either in person or over the phone should be given a

unique Subject Identification Number. The following process describes how to obtain a Subject

Identification Number.

1.

Log on to the study website using a computer with a secure Internet connection. The

location to obtain a study ID number is:

https://biosweb.njc.org/sec/COPDGene/register.cfm

[Alternately, you may log in to the study site home page and then choose “Get an ID” on

the left-hand side of the page. Use the login name and password provided to you by the

Data Coordinating Center (DCC).]

COPDGene®

Version5.0 June 15, 2009

– 17 –

Manual of Operations

Subject Enrollment

Access to the site is password-protected. The DCC is responsible for issuing individual

login names and passwords to COPDGene® personnel. If you forget this information you

may contact Ruthie Knowles (KnowlesR@njhealth.org, 303-270-2133) at the DCC.

The subject ID consists of a 5-digit number followed by a character. The character is a

check-digit that catches transposition of numbers in the numeric portion of the ID. In order

to facilitate identification of the Center at which a subject is enrolled, a 3-character Clinical

Center prefix is attached automatically to the ID: for example, NJC_99999Z.

2.

3.

Fill in the information requested in order to receive a Subject Identification Number.

When an ID number is assigned, print the page with the ID number and keep this

information with the source documents at your site for future reference. A sample subject

identification form is available on the COPDGene® website.

Calculation of Age

4.

This section clarifies the calculation of age for a COPDGene® subject. This calculation is done

when a coordinator gets a subject ID at the following secure web site that is password protected:

https://biosweb.njc.org/sec/COPDGene/sm/register.cfm.

This page is reached by clicking “Get an ID” under the Subjects heading of the left-hand menu

of the COPDGene® web site. The highlighted text (see above) is certification by the

COPDGene® Data Coordinating Center that date of birth is used only to calculate age.

Depending upon the Clinical Center and their local IRB approval, age is rounded to the nearest

COPDGene®

Version5.0 June 15, 2009

– 18 –

Manual of Operations

Subject Enrollment

0.10 year or to 0.25 year (the latter for Veterans Administration Centers) on the computer of the

user.

When a user goes to the web page

https://biosweb.njc.org/sec/COPDGene/sm/register.cfm

to get a subject ID, the web page itself, the coding that generates the web page, and the coding

that calculates age are loaded within the browser (e.g., Internet Explorer, Firefox, Safari) that is

running on the computer of the user. When date of birth is entered on this web page, that

information goes no farther than the browser that resides on the computer of the user. Once the

date of birth is calculated (rounded to the appropriate age per local IRB approval), the date of

birth is cleared from the screen and no longer resides on the computer.

In summary, date of birth is only entered at the COPDGene® Center that is obtaining a subject ID

and once the age is calculated the date of birth is not retained on the local computer.

Date of birth is also entered on the Demographics and Physical Characteristics Form. Date of

birth is only entered on this form as allowed by the Clinical Center IRB approval and HIPPA

restrictions. The date of birth entered on this form is only used as a check of the age calculated

on the Get a Subject ID form.

4d.

Barcode Labels

Once a subject receives a Subject Identification Number, barcode labels will be emailed to the

coordinator. Barcode labels should be printed at the Clinical Center prior to the actual study visit

since they will be needed during the study visit.

These labels will have a preprinted COPDGene® ID number that looks like this:

The labels are formatted for Avery 5267 label sheets. When printing the labels, make sure in the

Page Handling section of the Print screen that Page Scaling is set to “None” before you print the

labels. If you have questions about or problems with the labels, contact Ruthie Knowles

(KnowlesR@njhealth.org, 303-270-2133) at the DCC.

Barcode labels should be attached to the following:

• All blood tubes

• Blood Collection Form

• CT Acquisition Form

• Chest CT scan DVD holder

• CT Shipment Form

• Chest CT Assessment Form

• Subject’s Study Folder

COPDGene®

Version5.0 June 15, 2009

– 19 –

Manual of Operations

Subject Enrollment

4e.

Informed Consent Administration

The study should be explained to potential subjects who pass the Preliminary Eligibility

Assessment (if IRB approved at the Clinical Center to permit use of this form before informed

consent is obtained). Subjects should be allowed to ask questions and may request additional

time to review the informed consent and study with family members prior to providing informed

consent. The informed consent form should be reviewed thoroughly with the subject. The

coordinator must assure that the subject understands the study and informed consent.

4f.

Eligibility Assessment

Prior to the Study Visit: Make sure you have obtained a Subject Identification Number. A

Preliminary Eligibility Assessment may be completed prior to the Study Visit, if approved by the

Clinical Center IRB. This may be done at the first contact with a potential subject by phone or in

person. Use of this Form will help assure that potential subjects are not unnecessarily scheduled

for a Study Visit. Check on the “Preliminary” button on the Form only when entering

information for screening prior to an in-person Study Visit. This information may be saved for

later review and should also be printed as a source document. To save the preliminary

information, check the box to confirm all responses, and click on “Save Preliminary”.

If preliminary data is entered, this information can be retrieved at a later time for that subject, as

long as it has been saved. To retrieve the Preliminary Eligibility Assessment, re-enter the subject

ID, select “Preliminary screening”, and click on “Edit Preliminary”. The data entered previously

for this subject will be shown on the screen.

At the In-Person Study Visit: A Final Eligibility Assessment must be performed to assure the

subject is still eligible and that information has not changed since the time of the Preliminary

Eligibility Assessment.

When completing a Final Eligibility Assessment on the website, the study coordinator should

perform this assessment based on information obtained during the face-to-face Study Visit from

the potential study subject. This step is required before any questionnaires are administered,

before spirometry, and prior to any testing. This form incorporates the inclusion criteria and

exclusion criteria for the study and assures subjects meet these criteria prior to any evaluations.

Subjects who do not pass the Eligibility Assessment are considered to have failed screening.

Even if a Preliminary Eligibility Assessment was completed prior to the in-person Study Visit, a

Final Eligibility Assessment MUST be completed with the patient after the informed consent is

obtained to assure it is accurate and correct. For example, a subject may have had a COPD

exacerbation between the time of the Preliminary Eligibility Assessment and the in-person visit –

this would exclude the subject from study procedures. (Note: Subjects who have had a recent

exacerbation may be enrolled at a later date after the exacerbation is resolved – one month or

more after completing antibiotics or steroids prescribed for the exacerbation.) Do NOT click the

FINAL button until you are ready to reenter all of the eligibility information. Clicking on

FINAL will clear all data.

COPDGene®

Version5.0 June 15, 2009

– 20 –

Manual of Operations

Subject Enrollment

Follow this process for the Final Eligibility Assessment during the in-person Study Visit.

a.Enter the subject ID and click on “Final.”

b.Complete all fields.

c.At the bottom of the form, there is a checkbox to confirm that data has been

reviewed (to reduce input errors). Confirm you have entered the information

correctly and click on the “Confirm” button.

d.This will then bring up the Print and Submit buttons.

e.“Print” the form for your on-site records

f.“Submit” the form to the DCC.

INCLUSION/EXCLUSION CRITERIA

The inclusion/exclusion criteria for the two groups of subjects (COPD and non-COPD) are

available in the study protocol and are also listed here for convenience.

General Inclusion and Exclusion criteria for all study subjects

The following criteria will be required on ALL smoking and non-smoking study subjects:

Inclusion criteria

Age 45 - 80 years

Non-Hispanic Whites and African Americans

Exclusion criteria

Smoking , ten pack years of more than 100 cigarettes in a lifetime

Physician diagnosed respiratory disease other than COPD or asthma (based on subject

report)

Other concomitant respiratory disorder(such as, but not limited to, diffuse

bronchiectasis, cystic fibrosis, or interstitial lung disease)

Lung surgery with removal of a lobe or more (including lung volume reduction and

lung transplantation)

Lung cancer, known or suspected

Bronchoscopic lung volume reduction

Pregnancy or suspected pregnancy

Uncontrolled cancer, as defined as ongoing radiation therapy, ongoing chemotherapy,

narcotics for pain control, or known metastatic disease

History of radiation therapy to the chest (other than radiation for breast cancer)

Use of antibiotics and/or systemic steroids (new prescription or increased dose) for a

COPD exacerbation within the last month

Inability to use albuterol

First or second degree relative (parent, brother, sister, daughter, son, aunt, uncle,

nephew, niece, half-sibling, grandparent, grandchild) of a subject enrolled in

COPDGene®

Subjects who indicate they are in more than one racial category

Metal objects that may interfere with Chest CT quantification including presence of a

cardiac pacemaker, defibrillator, metal prosthetic heart valve, metal projectile or

metal weapon fragment (bullet, shrapnel, shotgun shot) or metal shoulder prosthesis

Subjects unable to perform spirometry due to:

COPDGene®

Version5.0 June 15, 2009

– 21 –

Manual of Operations

Subject Enrollment

• chest or abdominal surgery in the past three months

• a heart attack in the last three months

• detached retina or eye surgery in the past three months

• hospitalization for any other heart problem in the past month

Participation in the ECLIPSE study

Inability to provide two telephone contact numbers

Place of permanent residence of three months or more

• COPD Subjects:

Additional Inclusion criteria

Smoking history of > 10 pack-years

Diagnosis of COPD (post-bronchodilator FEV1/FVC < 0.70) Stages 1, 2, 3 and 4 by

GOLD criteria (7)

Additional Exclusion criteria

Smoking history of < ten pack years

Notes (These notes generally apply to both subjects with and without COPD):

Age: Exclusions to allow enrollment of subjects who do not meet age criteria will not be given.

Subjects must be 45 years of age or older at the time of the Study Visit and must be 80 years of

age or younger at the time of the Study Visit.

COPD Diagnosis: The diagnosis of COPD includes airflow limitation on spirometry and a

history of risk factors (most commonly cigarette smoking) known to cause COPD. We will use

NHANES (Hankinson) predicted spirometry values obtained in the United States. Subjects will

be placed in a COPD Stage after post-bronchodilator spirometry that demonstrates airflow

limitation (FEV1/FVC < 0.70) obtained during the Study Visit. Subjects in the COPD Group do

not need to have a prior physician’s diagnosis of COPD. Subjects who have been given a

diagnosis of COPD but who have normal spirometry will be enrolled as control subjects.

Cigarette Smoking: There are no uniform criteria for the amount of cigarette smoking required

for the diagnosis of COPD. We have chosen a threshold for cigarette consumption (10 pack-

years) to ensure that all subjects have a substantial environmental stress to differentiate those

who have an abnormal pulmonary response to cigarette smoke and those who do not have such a

response. Non-smoking controls will be defined as having smoked less than 100 cigarettes, 52

cigars or 12 oz. of pipe tobacco in a lifetime.

Other Diagnoses: Subjects will be included in the COPD group if they have evidence of airflow

limitation “that is not fully reversible” using the GOLD (Global Initiative on Obstructive Lung

Disease) criteria with a post-bronchodilator FEV1/FVC < 0.70. The distinction between asthma

and COPD is often difficult. Review of medical records of patients with COPD frequently

demonstrates a physician-listed diagnosis of asthma. Elimination of asthma patients with fixed

airflow limitation would inappropriately bias the findings of this study. Thus, if a subject meets

COPDGene®

Version5.0 June 15, 2009

– 22 –

Manual of Operations

Subject Enrollment

the smoking history and spirometry criteria for the study, they should NOT be excluded if they

have a concomitant diagnosis of asthma.

We will employ the FVC as the primary measure of lung volume since this will allow us to

compare the study results to previous large-scale epidemiologic investigations. Normal values

have been published, and this maneuver is widely used. Subjects who are known to have

interstitial lung disease are excluded.

COPDGene® is not targeting patients with known alpha-1 antitrypsin deficiency. Coordinators

should NOT recruit patients with known alpha-1 antitrypsin deficiency. However, if patients are

enrolled in the study and later found to have alpha-1 antitrypsin deficiency they will not be

excluded. Alpha-1 antitrypsin deficiency is a known genetic risk factor for COPD and enrollment

of large numbers of such patients will weaken our ability to find new genetic determinants of

COPD.

Prior Lung Surgery or Lung Volume Reduction: Subjects are excluded if they have had a lobe or

more of one lung removed, lung volume reduction surgery, bronchoscopic lung volume

reduction, or lung transplant. Such surgery will impact lung function such that it will be

impossible to determine the Stage of COPD.

Lung Cancer: If there is an excess number of subjects with lung cancer among COPD cases in

this study, then we may detect genes associated with lung cancer rather than genes associated

with COPD. Subjects with known or suspected lung cancer will be excluded. Subjects with a

prior history of lung cancer, even if successfully resected and cured will be excluded. Subjects

who may in the near future have a resection of a lesion suspected to be lung cancer will be

excluded; these subjects may be included at a later time if the lesion proves not to be cancer and

the resection removes less than one lobe of the lung. Subjects with small, indeterminate

pulmonary nodules are eligible to participate.

Uncontrolled Cancer: Subjects with uncontrolled cancer of any type will be excluded because the

cancer or cancer treatment may alter lung function and thus misclassify the respiratory status of

the subject. In addition, we expect to conduct long-term follow-up the cohort of subjects enrolled

in this study; subjects with uncontrolled cancers may expire within 5 years and prior to future

planned reassessment and follow-up. These principles and the details noted above in these

exclusion criteria should be used to determine which subjects should be excluded. The following

examples may be used as a guide. A potential subject with a history of prostate cancer that has

not been treated may be included in COPDGene®. A potential subject with a history of breast

cancer without known metastases who had surgery two years ago may be included in

COPDGene®.

Radiation Therapy: Subjects who have had radiation therapy to the chest or mediastinum should

be excluded. Radiation therapy may alter lung function making it difficult to assess the presence

and Stage of COPD. Subjects who have had radiation therapy for breast cancer may be included

in COPDGene® since modern radiation techniques do not usually cause radiation changes in the

lungs.

COPDGene®

Version5.0 June 15, 2009

– 23 –

Manual of Operations

Subject Enrollment

Recent Antibiotic or Steroid Use: A recent COPD exacerbation may temporarily alter lung

function thus misclassifying the presence of COPD and the COPD Stage of such subjects. In

addition, a recent COPD exacerbation may have an ongoing inflammatory pulmonary process or

an infection that will impair the ability to diagnose, characterize and phenotype subjects in their

usual stable state. Subjects who have had antibiotics or systemic steroids for a COPD

exacerbation can be re-screened 30 days after cessation of antibiotics or corticosteroids. Subjects

who have used an antibiotic for an infection other than a respiratory infection may be enrolled at

any time they are clinically stable and do not have to wait 30 days. An example would be

outpatient antibiotic treatment for a urinary tract infection.

Inability to Use Albuterol: Albuterol administration is required to obtain post-bronchodilator

spirometry to assess the presence and severity (Stage) of COPD. Subjects who refuse to have

albuterol administered should be excluded. Based on the Safety Assessment, subjects who have

had or have the potential to develop adverse effects of albuterol should be assessed by a

physician investigator to determine if it is safe for the subject to receive albuterol and continue in

the study.

Relative of a COPDGene® Subject: Subjects who are related to another subject already enrolled

by marriage are allowed to participate. For example, a sister-in-law of a COPDGene®

participant may be enrolled. A spouse of a COPDGene® subject may be enrolled.

More than one racial category: Subjects are excluded if they self-designate that they are in more

than one racial category.

Metal Objects in the Chest: Metal objects in the chest or chest wall will interfere with the chest

CT scan. This may make assessment of the amount of emphysema and airways impossible.

Metal objects in the chest include defibrillators, pacemakers, metal shoulder prosthesis and

prosthetic metal heart valves, metal projectile or metal weapon fragment (bullet, shrapnel,

shotgun shot). If there is a question about other metal objects, contact your local PI, COPDGene®

radiologist, or the COPDGene® Imaging Core.

• Smokers without COPD

Additional Inclusion criteria

History of cigarette smoking > 10 pack-years

Post-bronchodilator FEV1/FVC > 0.70 and FEV1 > 80% predicted(9).

Additional Exclusion criteria

Smoking history of < ten pack years

Notes:

Lung Disease: The primary criterion for determination of COPD cases vs. non-smokers without

COPD is spirometry, not physician diagnosis. However, subjects with pulmonary diseases

including, but not limited to, interstitial lung disease, diffuse bronchiectasis, or cystic fibrosis,

are not eligible for COPDGene®. Physician-diagnosed lung disease other than COPD is an

exclusion for non-COPD subjects.

COPDGene®

Version5.0 June 15, 2009

– 24 –

Manual of Operations

Subject Enrollment

Note: In order to avoid potential biases involved in allowing asthmatic subjects to be part of the

COPD case group but not the control group, asthmatic subjects are eligible to be either cases or

controls. However, an excess of asthmatic subjects in the control group could interfere with the

genetic association results. Therefore, Clinical Centers are encouraged not to enroll control

subjects from sources that are enriched for asthmatics (e.g. Asthma Clinics).

Subjects may have been previously told by a physician or other health-care practitioner that they

have COPD, emphysema or chronic bronchitis. Such subjects who have normal spirometry do

not meet spirometric criteria for COPD and will be included in COPDGene®. In such cases

where COPD is suspected, the Final Eligibility Form may have been submitted with a positive

response to the question about physician-diagnosed lung disease. If this form needs to be

changed based on normal spirometry results and the subject needs to thus be re-classified as not

having COPD, coordinators should first obtain confirmation from their Clinical Center

Investigator. Then the coordinator should contact Maura Robinson at the Data Coordinating

Center to inform her of the need to change the Final Eligibility Form.

Spirometry: Sources of subjects will vary from center to center, but will likely include inpatient

and outpatients at the centers, spouses and friends of subjects with COPD, patients in primary

care practices, local patient support and educational groups, and local and national COPD

voluntary organizations (such as COPD Foundation and American Lung Association).

Subjects should not primarily be recruited from sources that include a high prevalence of

asthmatics such as asthma clinics or asthma patient groups. While subjects with asthma are not

excluded in either control or COPD subjects in order to assure similar inclusion/exclusion

criteria in both populations, the study is not designed as a study of the genetics of asthma. Thus

subjects who have asthma as their primary respiratory disease should not be targeted for

recruitment.

Although the primary focus of this project is COPD and COPD-related phenotypes, subjects will

also be informed that this cohort may be used to study the genetic and environmental determi-

nants of other smoking-related illnesses such as lung cancer and coronary artery disease and,

with their permission on the consent form, other disorders that are not smoking-related.

• GOLD Unclassified Subjects:

Inclusion criteria

Smoking history of ≥ 10 pack-years

Spirometry (Post-bronchodilator FEV1/FVC > 0.70 FEV1 <80% predicted)

Exclusion criteria

Smoking history of < ten pack years

Subjects that meet the required smoking history parameter but do not fall into either COPD or

smoking control categories will be included as Smoking Unclassified subjects. These subjects

will be grouped based on the occurrence of a normal spirometry accompanied by a presence of

reduced air flow (post-bronchodilator FEV1/FVC > 0.70, lowFEV1 <80%). To date, we have

found approximately 10 - 12% of our smoking subjects fall into this category. The inclusion of

COPDGene®

Version5.0 June 15, 2009

– 25 –

Manual of Operations

Subject Enrollment

the Unclassified smoking group will provide additional and possibly novel information regarding

the categorization, clinical presentation, and progression of COPD in smoking populations. Since

this group is already present within the current subject pool, these subjects will be clarified as a

distinct subject group.

• Non-smoking Controls

Additional Inclusion criteria

No smoking history as defined by less than 100 cigarettes, 52 cigars or 12 oz. of pipe

tobacco smoked in a lifetime

No airflow limitation (Post-bronchodilator FEV1/FVC > 0.70 FEV1 > 80% predicted)

Additional Exclusion criteria

Smoking history of more than 100 cigarettes, 52 cigars or 12 oz. of pipe tobacco smoked in

a lifetime

No physician diagnosed history of respiratory disease

Subjects with no smoking history and no airflow limitation will be included as a reference

population for comparison with those affected by smoking exposure. These subjects will also

offer information on the processes involved in normal lung aging. Non-smoking control subjects

are critical for supporting genetic and pathological findings within smoking and diseased subject

groups by acting as a baseline for normal pulmonary physiology within a genetically mixed

population.

Race/Ethnicity

This information is required as part of the eligibility assessment. Subjects will be asked to

provide their ethnicity and race according to National Institutes of Health (NIH) categories.

Subjects should be shown the following table and

• first asked to identify one ethnic category and then

• indicate one or more racial categories that they believe most accurately describes

themselves.

Subjects will only be enrolled if they indicate a racial category of Black/African American or

White. Potential subjects who indicate more than one race may not be enrolled in COPDGene®.

Only subjects who self-indicate they are non-Hispanic should be enrolled.

COPDGene®

Version5.0 June 15, 2009

– 26 –

Manual of Operations

Subject Enrollment

ETHNIC CATEGORY (indicate only one)

Hispanic or Latino

Not Hispanic or Latino

RACIAL CATEGORY

Black or African American

White

American Indian / Alaska Native

Asian

Native Hawaiian or Other Pacific Islander

Subjects should be shown the chart above and asked to self-identify race and ethnicity. Subjects

may choose as many responses as they feel are appropriate. This should be completed by the

subject without coaching from the coordinator or investigator.

COPDGene®

Version5.0 June 15, 2009

– 27 –

Manual of Operations

Subject Enrollment

5.

Collection of Other Medical Information

Every reasonable effort should be made to collect other relevant medical information from the

subject’s primary care provider or physician specialists. The patient should sign a consent for

release of medical information to obtain protected medical information. The following

information should be obtained if available:

• High resolution CT scan within the last year if an appropriate CT protocol was used and if the

required computerized archive is available for computer analysis,

• Pulmonary function tests including lung volumes and diffusing capacity within the last year,

• Arterial blood gas within the last year.

COPDGene®

Version5.0 June 15, 2009

– 28 –

Manual of Operations

Other Medical Information

6.

Subject Enrollment and Evaluation Sequence

IMPORTANT NOTE:

In order to enroll a subject and perform a chest CT scan in COPDGene®, steps 1 through 14

below MUST be followed precisely in the order specified here.

The order of Steps 1 through 14 is to assure that subjects meet inclusion and exclusion criteria as

soon as possible and as efficiently as possible. Subjects may screen fail at any point during these

initial steps. All data must be collected within six months of initial patient visit. Chest CT scans

will be done on all subjects.

The order of these procedures during the Study Visit is mandatory:

1. Prior to the study visit, obtain a unique Subject Identification Number

2. Prior to the study visit, print barcode labels that are emailed from the COPDGene® Data

Coordinating Center to the principle coordinator

3. Perform Preliminary Eligibility Assessment over the phone or in-person, if IRB

approved at the Clinical Center

4. Obtain informed consent from the subject

5. Complete Informed Consent and Permissions Form

6. Perform FINAL Eligibility Assessment to determine if subject is eligible to continue

7. If a pregnancy test is required, you should do it at this point prior to spirometry and

albuterol.

8. Complete Demographics-Physical Exam Form including resting heart rate, oxygen

saturation, blood pressure, height and weight

9. Complete Safety Assessment Form

10. Perform spirometry before bronchodilator

11. Administer 2 puffs of albuterol HFA with an appropriate spacer such as an Aerochamber

12. Wait 15 - 20 minutes after albuterol. It is suggested the wait be as short as possible to

most efficiently determine if the subject will meet entry criteria. While you are waiting,

you may collect contact information during this period and/or administer brief

questionnaires.

13. Perform spirometry after bronchodilator and complete Spirometry Form

14. Complete Assignment Form to determine if the subject meets inclusion/exclusion

criteria. This form must be completed on a computer with an internet connection. This

form is necessary to determine if the subject should continue in the study for the rest of

the visit including the Chest CT scan.

If a subject fails to meet study entry criteria at any time during the procedures above a

Discontinuation Form must be completed. This form is also used if the subject withdraws

permission to participate at any time.

The target for post-bronchodilator spirometry is 15 - 20 minutes post albuterol administration.

The window for post-bronchodilator spirometry is 15 - 40 minutes post albuterol.

After the initial Steps 1-14 are completed, the remaining Questionnaires, Forms, Walk Test, CT

scan and blood draw must be performed. The remainder of the procedures may be performed in

the order most convenient for the subject, coordinator, Radiology Department and local Clinical

COPDGene®

Version 5.0, June 10 2009

Manual of Operations

Evaluation Sequence

– 29 –

Center operations. The order of the remaining procedures might also be based on these additional

considerations:

• Potential subject fatigue, especially those who are older or have more severe COPD.

• Subject ability to concentrate on questionnaires for long periods of time. Questionnaires

might be interspersed with other procedures like walk test and CT scan.

• Ability of subjects to walk for long distances to other procedures like spirometry (pre/post

bronchodilator testing), the chest CT scanner and six-minute walk course

• Potential subject fatigue after the Walk Test

• Potential subject fatigue after spirometry

• Scheduling of Chest CT scanner

Consideration should be given to assure that subjects are not overly taxed physically or mentally

and that sufficient time is given to allow a rest if walking between testing sites is required. For

example, subjects should not be required to walk a long distance to the six-minute walk test

course and then immediately perform the walk test. Rather, a rest may be indicated prior to

starting the walk test, at which time questionnaires may be administered. Alternatively, a

wheelchair can be used to transport the subject to the Walk Course. Subjects may become

mentally fatigued if required to spend a long time without breaks completing questionnaires.

Alternating questionnaires with other tests such as blood draw, walk test and spirometry may

avoid excessive mental fatigue. Coordinators should plan tests in a sequence to be as efficient as

possible and minimize patient time and coordinator effort. It is suggested that spirometry before

and after albuterol be performed with a questionnaire administration during the time between

albuterol administration and post-bronchodilator spirometry. Spirometry should not be

performed immediately following blood sampling. The Walk Test should be done at least 20

minutes after the albuterol administration and after the post-bronchodilator spirometry.

As an example, after the initial Steps 1-14 above, one test sequence of the remaining

questionnaires and procedures that may maximize efficiency is:

15. St. George’s Respiratory Questionnaire, SGRQ (to allow a rest after spirometry)

16. Respiratory History Questionnaire

17. MOS SF-36 Questionnaire

18. Blood Sampling

19. Medical History Questionnaire

20. Medication Questionnaire

21. Six-minute Walk Test (If the distance the subject must walk to get to the walk course is

lengthy, then walking to the course and resting prior to starting the test may be helpful.

During the resting time, the medical history and medication questionnaires may be

administered.)

22. Chest CT Scan

COPDGene®

Version 5.0, June 10 2009

– 30 –

Manual of Operations

Evaluation Sequence

7.

Questionnaires

Questionnaires can be completed on a computer copy of the questionnaire or on a printed copy of

the questionnaire.

• If you are completing Forms and questionnaires using a printed copy, you should print the

forms and place the Subject Identification Number on each form prior to the Study Visit.

After the questionnaires are completed, they must be entered into a computer copy of the

Form and transmitted to the Data Coordinating Center. Double-entry of data is not

necessary, but quality control of entry of printed forms will be performed periodically.

Maintain the original form completed by the subject as a Source Document. Maintain a

computer copy of each form as well on the study computer at your site.

• If you are completing Forms and questionnaires on the computer, create a folder on the

computer with the Subject Identification Number. Place a copy of each Form in the folder

and label each Form with name of the Form and the Subject Identification Number. Use a

name for the Forms in a consistent manner that is most convenient for you and allows you

to identify the Subject and the Form in the future; for example you might re-name the

COPDGene® Spirometry Form as “NJ_99999Z Spirometry.”

Materials for Administering Questionnaires:

• Computer with Internet connection (at least 15 inch screen suggested if the subject is

completing the questionnaire on the computer)

• Adobe Acrobat (Standard or Professional Edition) installed on computer. This is not a free

program; rather it must be purchased for each computer used for the study

• Downloaded questionnaires from the COPDGene® website.

Study questionnaires and forms are in Adobe Acrobat™ pdf format and require purchase and

installation of a full version of Adobe Acrobat™ on each computer that will be used for subject

interviews. For more information see the Adobe website:

http://www.adobe.com/products/acrobat/.

Questionnaires are downloaded from the study website:

https://biosweb.njc.org/sec/COPDGene/MainPage.cfm.

A new questionnaire should be used for each subject and completed on a computer. The

coordinator must administer each questionnaire and fill in the subject’s responses on the

computer form. The coordinator MUST review all completed questionnaires with the subject to

assure that all questions are answered by the subject.

Important Note for Administering Questionnaires:

All questionnaires must be administered as objectively as possible. The coordinator must keep

“clinical judgment” out of assessing the subject’s responses. Use the exact wording of each

question. If the subject expresses doubt as to the meaning of the question, repeat it exactly.

Emphasizing individual words or phrases often makes the meaning clear. Further explanation

may be needed, but do not cross-examine the respondent. When, after brief explanation, doubt

remains as to whether the answer should be “yes” or “no,” the answer should be recorded as

“no.”

COPDGene®

Version5.0 June 15, 2009

– 31 –

Manual of Operations

Questionnaires

7a.

Demographic Data and Contact Information

A Microsoft Excel spreadsheet is provided on the study website to facilitate collection of

demographic information at the Clinical Center. This information should be kept at the Clinical

Center and is NOT transmitted to the DCC. This information will be used by the site to contact

the subject for future studies twice a year. This contact may be in the form of a phone call, a

newsletter with a reply card, and/or an email with a required response.

All subjects must have the following demographic information collected and stored at the

Clinical Center and maintained in a secure manner:

• name,

• social security number

• permanent address,

• email address,

• two telephone numbers

• date of birth,

• primary care physician name, address and phone number.

In order to facilitate contact with the subject for longitudinal data collection and vital status, the

following information must also be collected:

• For a close relative:

o name,

o address,

o home telephone,

o cell phone,

o email address (if available).

• For a second close relative not living with the subject:

o name,

o address,

o home telephone,

o cell phone,

o email address (if available).

7b.

Safety Assessment Form

Prior to administration of albuterol and prior to the walk test, the Safety Assessment question-

naire should be administered. This Form provides guidelines for when an investigator should be

contacted prior to performance of spirometry, administration of albuterol and performance of the

Walk Test. Coordinators may ask additional questions not on this form in order to assist the local

physician investigator in interpreting the subject’s answers. For example, there is a question on

this form about chest pain. In assessing chest pain, if the subject knows their chest pain is based

on non-cardiac causes, the coordinator collect this information and transmit such information to

the physician investigator who will make a final determination of whether the subject may

proceed with COPDGene® testing. Based upon the Clinical Center Physician Investigator review

of the Safety Assessment Form and possible interview with the potential subject, a decision

should be made about the safety of continuing with the COPDGene® procedures.

COPDGene®

Version5.0 June 15, 2009

– 32 –

Manual of Operations

Questionnaires

7c.

Medical History Questionnaire

A medical history questionnaire will be administered to all subjects. This will be used to obtain a

history of other physician-diagnosed medical conditions.

7d.

Medications Questionnaire

All medications prescribed by a health care provider (physician, nurse practitioner, physician

assistant) must be captured on this questionnaire. Remember that oxygen is a medication that

should be included on this form even if the subject does not bring it to the Study Visit.

The easiest way to complete this form is to start at the bottom and write all the medications

prescribed by a health care provider that are currently being taken by the subject.

For all prescribed non-respiratory medications, the coordinator should write the medication in

the text field. Either brand or generic names may be used. Only the medication should be

recorded; doses and frequency of daily use are not to be entered. Include both medications taken

on an as needed basis (prn) and medications taken on a regular schedule. If the subject is using a

combination medication (e.g. Advair™, Symbicort™, Combivent™), answer affirmatively to the

questions regarding each component of the medication as well as the combination itself.

After the Study Visit, the coordinator should then complete the questions regarding classes of

respiratory medications using the list of all medications recorded during the Study Visit. For

respiratory medications, classes of medications are recorded; the listed drugs are examples of the

class. If there are questions about a medication that is not listed, the coordinator should consult

the local investigator to confirm the correct class of respiratory medication. If a question remains

about a drug the investigator or Clinical Center Coordinator should contact the Clinical Center

Director, Barry Make, MD (makeb@njhealth.org; 303-398-1720).

CLASS OF MEDICATION

Short-acting beta-2 agonist

Anticholinergic inhaler

Oral beta-2 agonist

Oral theophylline

Long-acting beta-2 agonist (including

combination preparations that contain

a long-acting beta agonist

bronchodilator)

Long-acting anticholinergic

EXAMPLES

albuterol, salbutamol, pirbuterol (Maxair™),

terbutaline

Atrovent™, ipratropium

Albuterol

Theodur™

Salmeterol (Serevent™), Advair™,

formoterol (Foradil™, Performomist™),

Symbicort™

Spiriva™, tiotropium

COPDGene®

Version5.0 June 15, 2009

– 33 –

Manual of Operations

Questionnaires

7e.

Respiratory History Questionnaire

These questions have been taken from the ATS Adult Respiratory Questionnaire and are used to

gather a respiratory history on all subjects. Some of the questions are from the new version of the

ATS questionnaire. If interviewer administration is used (which is preferred), the coordinator

should ask the subject the questions and mark the responses on the computer form.

General Guidelines for Administration of the Questionnaire

• The questionnaire can be administered by interviewers or reviewed by interviewers in the

presence of the participant who completed the questionnaire. The Clinical Center

Coordinator should review each subject’s questionnaire individually to obtain

consistency.

• Always ask each question the same way every time you ask it. Read each question slowly

and clearly. You can read the question again if the subject says that they do not

understand you.

• Do not bias the subject with your comments. Do not interpret the question for the subject.

For clarification, you should read the question again slowly and refer to the specific

guidelines below.

• Always read the question as it is written on the questionnaire. Do not shorten or add

anything to the question. When asking the same questionnaire repeatedly, you might be

tempted to make your own minor modifications - please do not.

• Some questions may seem repetitive; they are asked for quality control. Instruct the

subject to answer the question even if it seems similar to a question they might have

already answered.

• Do not give “uncertain” as an option—you would like them to answer YES or NO. If

they truly cannot answer the question, then put “Don’t Know” where applicable or if

there is no area to indicate a “don’t know” then answer “NO.”

• Fill out the questionnaire carefully and legibly. Make sure all the questions have been

asked and that the responses are consistent.

• Check over the questionnaire before the subject leaves.

• If you need to make a change upon reviewing a printed copy of the questionnaire, cross

out the incorrect response, circle the correct response, and write your initials next to that

questionnaire item. Save the original hard copy of all questionnaires as Source

Documents.

Respiratory Symptoms

Question 2 - “On getting up” may be at night for those who work at night.

Question 4 - Phlegm. Emphasis should be placed upon phlegm as coming up from the

chest; postnasal drip does not count as phlegm. This may be clarified with the follow-up

question: “Do you raise it up from your lungs, or do you merely clear it from your throat?”

Some subjects admit to phlegm without admitting to cough. This claim can be accepted

without changing the replies to “cough”; however, if they indicate that they cough up

phlegm after responding negatively to the cough questions, the cough questions can be

reviewed with the subject. Phlegm coughed up from the chest but swallowed counts as a

COPDGene®

Version5.0 June 15, 2009

– 34 –

Manual of Operations

Questionnaires

positive response. Include, if volunteered, phlegm with first smoke or “on first going out-

of-doors.” “On getting up” may be at night for night workers.

Question 7 – Periods of cough and phlegm. This question is to identify persons with

exacerbations (periodic worsening) of their symptoms that are beyond the normal daily

variation for the individual subject. Wording on this question is reliant on the previous two

questions whereas the answers each individual gives will determine whether you include

“increased” in question phrasing. If a person answers “Yes” to a usual cough AND

phlegm, then include “increased” in the wording of the question. However, if the subject

had said “No” to the previous questions about cough and phlegm, only ask … “periods or

episodes of cough and phlegm…” omitting the “increased.”

Question 8 – Wheezing. Subjects may confuse wheezing with snoring or bubbling sounds

in the chest. A demonstration “wheeze” may help if the subject requests further

clarification.

Question 15 – Unable to walk because of another condition. In addition to asking this

question, you may OBSERVE if the subject is disabled from walking (e.g., in wheelchair,

amputated leg, etc.). If subject is disabled from walking by any condition other than heart

or lung disease, then fill in the reason at this item and skip to next section on COPD

Exacerbations in the Last Year.

Question 16 – No attempt is to be made to separate out breathlessness from cardiac

conditions.

Severity of COPD Exacerbations in the Last Year

Complete this section even if the subject responded “no” to the questions in the last section

on COPD Exacerbations in the Last Year. If the subject responds “yes” to the first question

in this section, go back and ask the subject to reconsider responses to the questions in the

last section on COPD Exacerbations in the Last Year.

Respiratory Conditions

Question 5 through 7 – Chronic Bronchitis, Emphysema, COPD. We are interested if

they have been labeled with the diagnosis of chronic bronchitis, emphysema or COPD,

rather than giving them a clinical definition or basing the responses from the section on

current respiratory symptoms.

Cigarette Smoking

Question 5 – Average cigarettes per day. Record a number. In order to clarify the concept

of average (mean number of cigarettes smoked over period from age first started to age

finally stopped smoking), you may state “Recognizing that sometimes you smoked more

and sometimes you smoked less, on the average of the entire time you smoked, how many

cigarettes did you smoke per day?” If the subject is a non-smoking control, ask if the

subject has smoked less than 100 cigarettes, 52 cigars or 12 oz. of pipe tobacco in his/her

COPDGene®

Version5.0 June 15, 2009

– 35 –

Manual of Operations

Questionnaires

lifetime. If the subject is not sure, ask them to give you a general idea of their smoking

habits. Asses whether the subject has smoked recently or only a few times in the past.

Pipe Smoking

Question 1. A YES answer requires smoking more than 12 oz of tobacco in a lifetime (or

8 pouches of tobacco, where one pouch equals 1.5 oz).

Educational and Occupational History

Questions 4 and 9 - Do not include “housewife” or “house-husband” as affirmative

occupation response and do not write them in the occupation field.

Family History

Question 1 - Ask all questions about father first, then all questions about mother.

7f.

St. George’s Respiratory Questionnaire

This health-related quality of life measure was developed specifically for patients with chronic

respiratory disease. The coordinator can ask the subject the questions and note the responses on

the computer form. If self-completion by the participant is performed, the research coordinator

should check to be sure all questions are answered; subjects should be asked to provide answers

to all questions. If the questionnaire is administered as a hard printed copy, the original should be

kept as a source document.

7g.

MOS SF-36 Questionnaire

The SF-36 is a widely used and validated quality of life instrument that will reliably assess

quality of life in both the non-smoking controls and the existing groups. It will provide

complementary information to the St George instrument for the COPD subjects and will allow us

to compare COPD subjects to other diseases. The MOS SF-36 can be completed by the average

subject in 5 minutes and consists of 36 questions reflecting 8 domains of health related quality of

life. We are now enrolling non-smoking controls who are not likely to have respiratory

symptoms and thus will have enough missing data on the St George Respiratory Questionnaire

(SGRQ) that we may not be able to get valid scores of their health related quality of life (HR

QoL). HR QoL is associated with severity of illness in COPD.

COPDGene®

Version5.0 June 15, 2009

– 36 –

Manual of Operations

Questionnaires

8.

8a.

Physical Characteristics

Height

Height and weight should be measured prior to performance of spirometry.

Measure the subject’s height either in centimeters (cm) to the nearest cm, or in inches (in) to

nearest 1/2 inch while the subject is standing erect, bare-footed or in stocking feet with their head

looking straight ahead. A stadiometer is the recommended device for height measurement.

If the subject is unable to stand, or has marked spinal deformity (e.g., severe kyphoscoliosis), an

arm span measurement may be used to estimate standing height. Have the subject stretch the

arms in opposite directions and attain the maximal distance between the tips of the middle

fingers. For white men, height = arm span/1.03, for African-American men height = arm

span/1.06, and for women, height = arm span/1.01. Be sure to denote the correct units of

measurement on the Physical Characteristics Form.

8b.

Weight

The subject’s weight should be measured using an accurate scale with the subject barefooted or

in stocking feet. Weight is required for body mass index determination for the calculation of the

BODE index, but is not required for most reference values for spirometry. Weight may be

recorded in kilograms (kg) to nearest 0.5 kg, or in pounds (lb) to the nearest pound. Be sure to

denote the correct units of measurement on the Demographics and Physical Characteristics

Form.

8c.

Pulse Oximetry and Heart Rate

Pulse oximetry should be obtained prior to spirometry and the Six-Minute Walk Test. Pulse

oximetry should be obtained with the subject in the seated position. Pulse oximetry will be

recorded after the subject has remained at rest in the seated position for at least five minutes. The

pulse oximeter will be placed on a finger without nail polish and will be considered valid only if

a strong pulse is demonstrable. The apparent median value obtained while observing the monitor

over a one-minute observation period should be recorded.

If the subject is using oxygen, the pulse oximeter should be placed on the subject’s finger first.

Next, the subject’s oxygen should be discontinued while monitoring the oximeter for a period of

five minutes. If the pulse oximeter reading falls to 82% or less, oxygen will be replaced and a

reading of 82% will be recorded as the subject’s oximetry. The apparent median value obtained

while observing the monitor over a one-minute observation period should be recorded.

The heart rate obtained by the pulse oximeter should also be recorded. The apparent median

value obtained while observing the monitor over a one-minute observation period will be

recorded.

COPDGene®

Version 4.1, December 1, 2008

– 37 –

Manual of Operations

Physical Characteristics

9.

9a.

Spirometry

Overview

Spirometry is a test that measures function of the respiratory system. It is one of the simplest,

most effective tests available for the assessment of lung function. The spirometer registers the

amount of air a subject breathes and the rate at which the air moves. The most common

spirometric test requires that the subject take a full, deep breath and then exhale as forcefully as

possible. The subject's effort is called the forced expiratory maneuver and most commonly only

measures the amount and speed of air that is exhaled.

COPDGene® uses a spirometer (the ndd EasyOne™ Spirometer) that measures flow and volume

by ultra-sound transit time. This spirometer meets American Thoracic Society spirometry

standards.

Because the results of spirometry testing are used to determine the presence and severity of

COPD, the measurement must be performed according to strict standards by technicians who

have been properly trained and certified in how to conduct the maneuver. In addition, the

equipment must be in good operating order and the calibration regularly checked. All spirometric

maneuvers will be reviewed by a central reading laboratory (COPDGene® Pulmonary Function

Core) to assure optimal quality of the data and to provide ongoing feedback to the pulmonary

function technicians regarding the adequacy of the maneuvers.

9b.

Summary of Measures

The following measurements will be obtained through spirometry testing during the

COPDGene® clinic visits:

A.

Forced Vital Capacity (FVC) is the total volume of air, expressed in liters, exhaled in a

forced expiratory maneuver (the act of exhaling as hard and fast as possible after a

maximal inspiration). The FVC is useful for detecting restrictive disorder, since lower than

expected results may be a sign that the lungs cannot inflate as fully as normal. The FVC

may also be reduced in people with more severe COPD and other obstructive disorders.

Forced Expiratory Volume at One Second (FEV1) is the amount of air, expressed in

liters, that a person breathes out during the first second of a forced expiratory maneuver.

This is reduced in people with limitations such as COPD and asthma. The FEV1 may also

be reduced in patients with restrictive disorders.

The ratio of FEV1 to the FVC (FEV1/FVC) is the most sensitive and specific index of

airflow limitation measured by a spirometer. It is obtained by dividing the FEV1 by the

FVC, and is expressed as a percent (i.e., 100 x FEV1/FVC). Note that the FEV1/FVC ratio

is the ratio of the absolute FEV1 and absolute FVC in liters and is not usually expressed as

a percent of predicted.

Forced Expiratory Volume at Six Seconds (FEV6) is the amount of air that a person

breathes out during the first six seconds of a forced expiratory maneuver. Increasing

interest is being shown in the FEV6, and more particularly in the FEV1/FEV6 ratio, as an

alternative to the FEV1/FVC ratio. Use of the FEV6 may be easier to obtain, particularly

for patients with severe airflow limitation with long forced exhalation times.

Manual of Operations

Spirometry

B.

C.

D.

COPDGene®

Version5.0 June 15, 2009

– 38 –

E.

The ratio of FEV1 to the FEV6 (FEV1/FEV6) is an alternative to the FEV1/FVC ratio. A

secondary objective of COPDGene® is to evaluate the utility of the FEV1/FEV6 ratio,

particularly with respect to the assessment of COPD.

Setting for Spirometry Testing

9c.

Spirometry testing ideally should be performed in a private, temperature-controlled room. All of

the instruments necessary for the test should be in the room. The room should be well lit,

preferably with a window, and located in a quiet area. These conditions will improve the quality

and reproducibility of the results. For safety, the participant should be seated in a chair with no

wheels; preferably the chair should have arms.

For some sites, it may be that testing will be done in non-clinical settings. The ndd EasyOne™

Spirometer is portable and has been shown to perform well in the field. Nonetheless, this

document assumes that testing will be done in a centralized clinic facility. Sites planning to do

otherwise should develop corresponding local procedures and document them in their local

Manual of Procedures.

Clean mouthpieces (Spirettes™), nose-clips, and spacers should be available in the room, as

should be a container to collect used Spirettes™ and used spacers. A box of facial tissue paper,

paper plate or some type of container to place dentures on or in (if needed), and a trash can

should be placed close to the participant. It also may be helpful for a source of drinking water to

be nearby, as some subjects may get dry mouth as they are performing the maneuvers.

The ndd EasyOne™ Spirometer does not need calibration. However, a calibration check should

be carried out daily to ensure that the spirometer is reading accurately. Instructions for

performing the calibration check are in the ndd EasyGuide™ technical manual and appear in

section 9n. The calibration syringe and adapter should always be stored next to the spirometer so

that the temperature between the syringe and the spirometer are the same. This will avoid having

large differences between room temperature and the spirometer temperature that could affect the

results of the calibration tests. If there is the potential for a large temperature difference between

the calibration syringe and the spirometer, the technician should pull and push the piston on the

syringe several times to correct the problem. If spirometry is done in the field (outside of a

clinical center), it is preferable to keep the spirometer and calibration syringe together overnight

to avoid temperature differences at the time of calibration.

9d.

Setting up the ndd EasyOne™ Spirometer

Prior to conducting the spirometry measurement, the technician should assure that the

configurations of the EasyOne™ are set according to the specifications outlined under the “ndd

Configuration Settings for Use in COPDGene®” below in sections 9j and 9p. Failure to have the

correct settings may result in lost or deleted maneuvers, which will result in the participant being

excluded from analysis in COPDGene®. Ideally, a single person should be designated as

responsible for configuration of the EasyOne™ at each Clinical Center.

COPDGene®

Version5.0 June 15, 2009

– 39 –

Manual of Operations

Spirometry

9e.

Setting up the Computer to be Used with the Spirometer

Install the EasyWareTM software from the CD included with the ndd spirometer on all PC

computers that will be used for spirometry the study. Note that a Macintosh computer can NOT

be used with the spirometer. It is strongly recommended that the spirometer be connected to a PC

with software installed during performance of spirometry in COPDGeneTM.

After installing the EasyWare™ software on your computer, follow the directions below to

change the default settings:

1. Open the EasyWare™ program on your computer.

2. Open the “File” pull-down menu.

3. Select “Preferences”

4. Mark the “Screen Connector” button (NOT USB Cradle or Serial Cradle)

5. Click “OK.”

9f.

Medication Use Prior to Testing

The subject’s recent bronchodilator use needs to be recorded on the COPDGene® Spirometry

Form. Some commonly used currently available bronchodilators and their classes are listed in the

table in section 7d above. In COPDGene®, bronchodilators should not be withheld prior to the

Study Visit. The reason for NOT withholding bronchodilators is to allow for all of the procedures

to be completed in a single Study Visit at which time the informed consent will be signed

immediately followed by all study procedures. While the assessment of bronchial hyper-

responsiveness may be biased because bronchodilators are not withheld, the post-bronchodilator

measurements should be unaffected, and it is the post-bronchodilator measurements that are used

to define and assess the Stage of COPD.

9g.

Bronchodilator Administration

In order to provide an assessment of bronchial hyper-responsiveness and to establish a diagnosis

of COPD and the Stage of COPD, a “Post” bronchodilator spirometry will be performed in

addition to the “Pre” bronchodilator test. In COPDGeneTM study, only albuterol will be used as

the bronchodilator. The brand names of some currently available albuterol HFA formulations are

ProAirTM and Ventolin HFATM.

9h.

Contraindications

The “Demographics and Physical Characteristics” and “Safety Assessment” forms should be

completed prior to spirometry to assure that administration of the test or the bronchodilator does

not pose a potential health risk. Since COPDGene® requires post bronchodilator assessment for

COPD determination, if the participant is unwilling or unable to provide a post bronchodilator

measurement, the subject will be excluded from the study. Specifically, spirometry testing

should not be done if the subject has or reports any of the following:

• chest or abdominal surgery in the past three months

• a heart attack in the last three months

• detached retina or eye surgery in the past three months

COPDGene®

Version5.0 June 15, 2009

– 40 –

Manual of Operations

Spirometry

• hospitalization for any other heart problem in the past month

• a resting pulse rate more than 120 beats/minute (participant should be sitting for at least

5 minutes prior to pulse rate determination)

In addition, if the participant exhibits any other co-morbidity (such as unstable angina or

pneumonia) that, in the opinion of a site clinician, may affect the performance of the test or

jeopardize the participant’s safety, then spirometry testing should not be done. Indicate this on

the Safety Assessment Form.

Note that the presence of a respiratory tract infection treated with an antibiotic in the four weeks

prior to the visit is a contraindication to testing in the COPDGene® study; this is an issue of not

only infection control but also for accurate diagnosis and Staging of COPD.

Ideally, sites could reschedule testing at a later date when the above situations are resolved. If

participants are brought back later for spirometry testing, the site should contact the Data

Coordinating Center for instructions on processing the data.

9i.

Conducting the Visit

A detailed description of the use and operation of the ndd EasyOne™ spirometer, together with

instructions for coaching the participant, are included in the ndd EasyGuide™ users’ manual. See

the ndd website (www.ndd.ch) for current version of users’ manual. All pulmonary function

technicians are expected to have read this document and to be familiar with its contents. A copy

of this document should be kept with each spirometer in case questions about the use of the

EasyOne™ spirometer arise during testing.

9i.1. Safety checks

Complete the Demographic Physical Exam Form and Safety Form prior to testing. Contact the

Clinical Center investigator to review the results if prompted to do so on the Safety Form.

For safety purposes, check each participant’s pulse rate prior to spirometry testing. Resting pulse

rate is determined by having the participant sit and rest for 5 minutes prior to the measurement.

Testing should not be done on subjects whose resting pulse is more than 120 beats per a minute

without prior approval from a physician investigator.

Coordinators should ask the safety questions that are included on the Safety Assessment Form. If

required by the Safety Assessment form, the PI will determine if the subject should continue with

Albuterol administration or not. Clearly document actions taken on the bottom of the Safety

Assessment Form.

9i.2. Preparing to conduct the spirometry maneuver

In order to minimize the risk of cross-contamination, technicians should wash their hands before

the start of the test and should use the Spirette™’s packaging to remove a mouthpiece (the

Spirette™) from its storage container for the participant to use. Insert the clean Spirette™ into

COPDGene®

Version5.0 June 15, 2009

– 41 –

Manual of Operations

Spirometry

the spirometer. Be careful to ensure that the arrow on the Spirette™ is lined up with the arrow on

the spirometer.

Testing should be conducted with the participant in the sitting position. A chair with arms and

without wheels should be used for the testing, and the participant should sit erect with chin

slightly elevated. The purpose of the chair is to support the participant in case she/he faints

during the maneuver.

Instruct participant to loosen any tight clothing that might restrict maximal inspiration. Some

individuals may have difficulty doing the spirometry maneuver due to urinary incontinence.

Offer the participant a chance to use the bathroom prior to testing.

Dentures should be left in place if they are not loose to help keep a tight seal around the

mouthpiece. If they are loose, have participant remove them and provide a paper plate or clean

container to place them on or in.

All maneuvers should be performed with the participant wearing a nose clip. This clip prevents

air from moving through the nose during the test. Noseclips should be cleaned with alcohol

wipes between tests or disposed of after each use.

9i.3. Enter the participant data into the spirometer

Prior to conducting the test, participant information must be entered into the EasyOne™. It is

easiest to do this via the EasyWare™ software on your computer (the spirometer must be

connected to the computer while you do this). Choose the “Edit” menu and the “New Patient”

option. Below are instructions for each field:

• ID field: Enter the participant’s COPDGene® ID number. This must be done in order to

link the participant’s questionnaire data with their spirometry data.

• Name field: Leave blank. Please do NOT enter the participant’s name into the ndd.

• Tech ID: Enter the 3-letter initials of the coordinator performing the test.

• Other fields: Enter the subject’s age, height, weight, ethnic category, gender.

Failure to enter all these data will result in incomplete data that may not be usable in

COPDGene®.

9i.4. Explain the purpose of the spirometry test

Explain that the purpose of the test is to check on the health of the participant’s lungs. Emphasize

that, although the procedure does not hurt, in order to get useful and valid results he/she must

breathe as hard and as fast as possible when told to do so and will need to repeat the procedure a

few times. Depending on the cultural setting in which the testing is done, subjects may need

repeated assurances that spirometry does not hurt them or damage anything.

COPDGene®

Version5.0 June 15, 2009

– 42 –

Manual of Operations

Spirometry

9i.5. Demonstrate the maneuver

Explain that the participant should:

- Take in as deep a breath as possible

- When his/her lungs are totally full, quickly position the mouthpiece

- BLAST out the air as hard and as fast as possible.

A vigorous demonstration of the maneuver by the technician will help to prevent wasted time

and effort caused by the participant’s lack of understanding. Demonstrate the correct placement

of the mouthpiece. The technician should take a deep breath and emphasize the maximal depth of

inhalation. He/she should then demonstrate the proper positioning of a demonstration Spirette™

mouthpiece and dramatically blast the air out as fast as possible.

9i.6. Pre-bronchodilator test

After instructing the participant about the procedure for spirometry testing, proceed with the

actual testing, following the procedures outlined in sections 5.2 to 5.4 of the ndd EasyGuide™

Users’ Manual and the detailed steps beginning on page 44.

The initial series of maneuvers is performed BEFORE administering the bronchodilator. Follow

the computer prompts until a successful test session has been obtained. A successful test session

is defined as at least three acceptable maneuvers, with the two best FEV1s and the two best FVCs

from these maneuvers both within 150 milliliters of each other.

9i.7. Administer the bronchodilator

After at least 3 acceptable and 2 reproducible maneuvers (see below for definitions of

“acceptable” and “reproducible”) are obtained, administer two puffs of bronchodilator (albuterol,

a short-acting beta-agonist) to the participant using a spacer. A large-volume spacer

(Aerochamber™ preferred) should be used. A timer should be set up to sound 15 minutes after

the last administered puff of bronchodilator. During the waiting time, the coordinator may

administer the study questionnaires. The effect of the bronchodilator will persist, and actually

slightly increase, for at least the next 30-40 minutes. Coordinators may therefore choose to fully

complete the questionnaires between the pre- and post-BD maneuvers, but should not wait more

than 40 minutes to do the post-BD spirometry.

9i.8. Post-bronchodilator maneuver

The post-bronchodilator (post-BD) maneuver can start anytime after the 15-minute wait. The

same criteria of at least 3 acceptable and 2 reproducible maneuvers should be followed. It is not

critical that the post-BD maneuver be done immediately at 15 minutes, but rather that it be done

at least 15 minutes but not more than 40 minutes after the last administered puff of

bronchodilator.

9i.9. Print results

Print the results for each subject and maintain a hard copy at the Clinical Center as a source

document.

COPDGene®

Version5.0 June 15, 2009

– 43 –

Manual of Operations

Spirometry

Clinical Centers wishing to give participants a hardcopy printout of their results can do so (if

IRB approved) following the instructions in the ndd EasyGuide™ Users’ Manual. Clinical

Centers choosing to give such printouts to their participants are encouraged to provide guidelines

for interpreting the results. It is strongly suggested that Clinical Centers provide reports to

subjects’ primary care physician (with subject consent and IRB approval).

9i.10. Coaching the participant and troubleshooting problems

Because the adequacy of these maneuvers is highly dependent on participant effort, the

technicians must guide the participant through the breathing maneuvers. It is extremely

important to inhale maximally and to exhale forcefully and maximally. Tell the participant when

to start taking in a deep breath and to put the mouthpiece in his/her mouth. Then tell the

participant to blast out the air and to continue exhaling for at least 6 seconds. Observe the body

language of the participant as he/she attempts to follow the instructions, and encourage the

participant to continue blowing out smoothly without re-breathing. Instruct the participant to

remain erect and not to bend over during the maneuver, and to keep their feet flat on the floor.

9j.

Step by Step Spirometry Instructions

a. Power on the EasyOne™.

b. Choose “Check Calibration” from the main menu on the spirometer.

c. Insert calibration Spirette™ into EasyOne™, matching arrows on both (designate one

Spirette™ for calibration use only).

d. Connect 3L syringe to Spirette™ using adapter.

e. Press ENTER on EasyOne™ to start the verification.

f. When tone is heard, withdraw one full pump stroke then inject one full pump stroke.

g. Ensure “Calibration Check Passed” message appears on EasyOne™ screen.

h. Choose “Quit.”

9j.2. Conduct the Pre-bronchodilator Spirometry

Note: The subject must have already completed and passed the Safety Assessment and if

necessary been approved for albuterol use by one of the site’s Investigators prior to spirometry.

a. Connect the EasyOne™ USB cable to the computer’s USB port.

b. Launch the EasyWare™ software via the desktop icon.

c. First time only: Choose “File”, then choose “Preferences,” and ensure that “Screen

Connector” is selected.

d. Power on the EasyOne™ and connect it to the cable. (“Device Connected” will display

at the bottom of the EasyWare™ window.)

e. From the Edit menu, choose New Patient, and enter the subject identification numbers,

three-letter Tech ID, date of birth, height, weight, gender, asthma status, ethnicity, and

smoking status. Confirm with OK.

f. Insert a new Spirette™ into the EasyOne™, matching the arrows on the front of both.

9j.1. Verify EasyOne™ Calibration

COPDGene®

Version5.0 June 15, 2009

– 44 –

Manual of Operations

Spirometry

g. From the View menu, choose Test On-Line.

h. To perform the Forced Vital Capacity maneuver, do the following on the EasyOne™

device keypad:

i. Use (>) to highlight ‘Recall’ and press ENTER to choose it

ii. Use ENTER to choose “Last Test”

iii. Verify the subject ID number and press ENTER.

iv. Press ENTER to choose FVC (Expiratory)

v. It is extremely important that the spirometer flow is zeroed. In order to zero

the spirometer, the EasyOne™ will prompt to “block the Spirette™”. At the

prompt,

(1) place the spirometer on a flat surface so that it is not moving,

(2) block the expired end of the Spirette™, and

(3) then press ENTER

vi. When prompted, with “Blast Out,” assist participant in applying nose clips and

instruct them to execute a forced expiratory maneuver.

vii. View results on computer. Review computer screen to assess adequacy of

spirometry. Reinstruct subject if necessary.

viii. Highlight “Retry” and press ENTER when ready to continue.

ix. Repeat steps vi through vii until three acceptable and reproducible efforts have

been recorded, in accordance with 2005 ATS criteria (best two FEV1 and FVC

volumes within 150 mL of one another), or until 8 efforts have been made.

Note: The EasyOne™ software may consider a session complete before

these criteria have been met. If that is the case, choose ADD to continue the

session.

x. If the “Acceptable” message appears, use (>) to highlight ‘Post’ and press

ENTER to choose it, preparing the device for post-bronchodilator efforts.

xi. If you do not see the “Acceptable” message, Use (>) to highlight ‘Quit’ and

press ENTER to choose it. Then Use (>) to highlight ‘Post’ and press ENTER to

choose it, preparing the device for post-bronchodilator efforts

xii. It is not necessary to include the pre-bronchodilator FVC and FEV1 on the

Spirometry form. However, you should retain a copy of all data at your site.

9j.3. Bronchodilator Administration

Once three acceptable, reproducible efforts have been made, use the spacer to administer two

puffs of Albuterol.

a. Instruct participant to exhale before each puff. Once dose is administered, instruct

participant to inhale slowly and deeply, and hold his/her breath for 10 seconds.

b. Record the time of bronchodilator administration on the Spirometry form. The post-

bronchodilator spirometry should be done 15-40 minutes later.

COPDGene®

Version5.0 June 15, 2009

– 45 –

Manual of Operations

Spirometry

9j.4. Post-bronchodilator Spirometry

a. Power on the EasyOne™ if it has shut itself down

b. To perform the Forced Vital Capacity maneuver, do the following on the EasyOne™

device keypad:

i. Choose ‘Perform Test’ with ENTER

ii. Use (>) to highlight ‘Recall’ and press ENTER to choose it

iii. Use ENTER to choose ‘Last Test’

iv. Verify that the info on the screen matches the participant you are testing

v. Press ENTER to choose this participant

vi. Ensure that the test says “FVC Post Med.”

vii. Choose ADD to add a post-bronchodilator session to this test.

viii. In order to zero the spirometer, the EasyOne™ will prompt to block the

Spirette™ – do so, and press ENTER

ix. When prompted, with “Blast Out,” assist participant in applying nose clips and

instruct them to execute a forced expiratory maneuver.

x. View results on computer. Review computer screen to assess adequacy of

spirometry. Reinstruct subject if necessary.

xi. Select RETRY when ready to continue.

xii. Choose Next to begin the next maneuver.

xiii. Repeat steps ix through xii until three acceptable and reproducible efforts have

been recorded, in accordance with 2005 ATS criteria (best two FEV1 and FVC

volumes within 150 mL of one another), or until 8 efforts have been made. NB:

The EasyOne™ software may consider a session complete before these

criteria have been met. If that is the case, choose ADD to continue the

session.

xiv. Quit when the session is complete.

xv. Look at the computer and complete the Spirometry form. Use the best values of

the post-bronchodilator FVC and FEV1 from the acceptable tests (see below).

The FVC and FEV1 do NOT need to be from the same test.

Complete the Assignment Form from the COPDGene® on the computer. Enter the post-

bronchodilator values for FVC and FEV1. Also enter the predicted values for FVC and FEV1

(the values are obtained from the pre-bronchodilator tests)—the percent predicted will be

calculated from these values.

9j.5. Print Your Results

a.

b.

c.

d.

On the computer in EasyWare™, highlight the row that contains the subject’s tests.

Go to file, select print.

Print the test.

Save a hard copy at the Clinical Center as a source document.

9j.6. Upload the Database File to the Website

The file that should be uploaded to the website is the Access database itself. If for some reason

you need to open the database and examine the tables, the password is “LungHealth”, but this

COPDGene®

Version5.0 June 15, 2009

– 46 –

Manual of Operations

Spirometry

should not be necessary. Further, if you open it, Access might try to “convert” the version. Never

choose to convert to a newer version – this will make the database malfunction with the

EasyWare™ software. Simply choose to open the database.

If you cannot find the location of this file on your computer, follow these steps:

1. Open the EasyWare™ software on your computer.

2. Click on “File”.

3. Click on “Preferences”.

4. In the box below “Database Path and Name”, is the place where your mdb file is located.

5. The default location for your database file is: C:\Program Files\ndd

Medizintechnik\EasyWare\EasyWare.mdb. You should find your file in this location,

unless you or someone else has changed the settings for the software.

All data that the PFT Core needs is in the "Easyware.mdb" file. Make a copy of this and upload

it to the COPDGene® website (see steps below). It would be helpful if you rename it when you

up load it. Please use the following format:

SITENAME_COPDGene_MM_DD_YYYY.mdb

For example: Uploading a file from Iowa on Feb 14, 2008, you would name the file the

following:

"IOWA_COPDGene_Feb_14_2008.mdb"

This will provide a unique name for every site on every day that they upload.

To upload the database to the website:

1. Open your internet web browser and go to the COPDGene® website

(https://biosweb.njc.org/sec/COPDGene/MainPage.cfm).

2. Log into the website. If you have lost your log-in information, contact Maura at the DCC.

3. On the left panel, click on “PFT Files.”

4. Log into the website again.

5. Click on thebutton across from the group called “COPDGene: Your Site.”

6. Click on the “Browse” button and locate your database (see the directions above if you

cannot find it).

7. Click on the .dmb file and choose to open it.

8. You may use the comments section to indicate if the new tests are for certification or

subjects.

9. Click “Add File Now.” You should now see your file in the table.

10. Once you add a file to a group, everyone in the group will be emailed that you have

uploaded a file to their group (you will not receive the email). These emails are sent from

the DCC hourly.

9k.

Acceptable and Reproducible Maneuvers

For the purpose of spirometry testing, “acceptable” is defined as a maneuver that is free from

error. “Reproducible” is defined as being without excessive variability between maneuvers. The

COPDGene®

Version5.0 June 15, 2009

Manual of Operations

Spirometry

– 47 –

following are some errors that can be seen or calculated from a forced expiratory maneuver and

that can affect acceptability: hesitation or false starts, cough, variable effort, glottis closure, early

termination, and leaks.

Three acceptable maneuvers are needed to determine reproducibility. The two highest values for

FVC and FEV1 taken from acceptable forced expiratory maneuvers must show minimal

variability (preferably within 150 milliliters of the second highest FVC and FEV1). It is also

important to inspect the volume-time curves to determine if the size and shapes of the curves are

reproducible.

The American Thoracic Society defines FEV1 and FVC as the best measurements from

acceptable and reproducible maneuvers. It is not necessary that they all come from the same

maneuver. The FEV1/FVC and FEV1/FEV6 ratios are computed as the ratio of the individual

measurements. In order to obtain these results, select the “best value” setting from the system

configuration menu (see section 8 of the EasyGuide™ Users’ Manual).

When errors occur, review common errors with the subject before proceeding with

additional maneuvers.

Ask the participant to watch the technician perform the FVC maneuver again. The technician

should demonstrate the correct placement of the mouthpiece, emphasize the maximum depth of

inhalation, and then blast out the air. If the participant tries again and the reproducibility criteria

are not met, the technician should continue administering the test as needed (up to a total of eight

maneuvers), assuming that the subject is able to continue.

Some participants will never be able to provide three acceptable and having two reproducible

maneuvers is OK. The goal is to meet the acceptability and reproducibility criteria, but these are

not absolute requirements for data to be used. Previous studies have shown that inability to

perform reproducible spirometry, even with good coaching, is an important risk factor in

predicting future health.

9l.

Reference Values

To interpret spirometric results, they must be compared either to a subject’s previous results or to

a published set of “predicted reference values.” Such predicted reference values typically

describe the average lung function for an individual of a given age, height, and sex, and such

equations have been published for a variety of racial groups.

Typically, lung function measurements are expressed as a percent of predicted. One hundred

percent of predicted represents the average value for the population, but normal, healthy

individuals will exhibit a wide range of values about this level. In general, values of 80% or

greater for FEV1 and FVC are considered to be in the normal range. Individuals whose test

results are below this level may have an abnormality.

The ndd EasyOne™ spirometer offers a number of published predicted values, most of which

were derived from studies of largely white participants. Four ethnic correction settings are

available that allow you to customize the amount of adjustment that is made for selected racial

COPDGene®

Version5.0 June 15, 2009

Manual of Operations

Spirometry

– 48 –

groupings. Consult the EasyGuide™ users’ manual, sections 8 and 12, for more information

regarding the use of prediction equations. For COPDGene®, we will use the reference values that

were derived from the 3rd National Health And Nutrition Examination Survey (NHANES III) in

the United States (Hankinson et al., AJRCCM 1999; 159:179-187). This study used a large,

randomly selected subset of the entire U.S. population, which included a variety of ethnic groups

and strict attention to quality control according to current American Thoracic Society guidelines.

Choice of reference values in the ndd EasyOne™ spirometry software only affects the printouts

available from the spirometry software, which sites may choose to give to the participants. For

the official COPDGene® dataset for each site, the COPDGene® Operations Center will create

local prediction equations based on healthy never smokers at that site. In addition, the official

COPDGene® dataset will contain additional prediction equations based on the United States

NHANES III equations for white participants.

9m. Implementation of Spirometry Quality Control

The Epidemiology Standardization Project, the American Thoracic Society spirometry standards,

and recent evaluations of commercially available spirometers emphasize the importance of

spirometry quality control (QC) procedures.

9m.1. Factors Affecting Spirometry Quality

a.Participant: A subject may not take as deep a breath as possible or exhale as forcefully as

possible at the start of the maneuver. Several possibilities will prevent a maneuver from being

acceptable: an involuntary epiglottis closure, temporarily cutting off the flow of air; an early

termination of the maneuver, preventing the achievement of a plateau; or a variable effort.

Coughing during the maneuver or a leakage due to the participant’s inability to keep a tight seal

also will prevent from obtaining a good maneuver. To address these sources of error, it is very

important to have technicians trained to watch the participant closely during the performance and

accurately review the displayed flow-volume curves on the computer monitor. The technician

thus can guide and explain the source of error to the participant.

b.Technician: Improper coaching or non-standardized coaching procedures will negatively

affect the quality of the spirometry testing. The technician should clearly instruct the participant

on how to perform this test, demonstrate the maneuver, and watch the participant closely during

the performance to avoid unacceptable errors and obtain the best effort from the participants.

He/she must be trained to recognize patterns of unacceptable maneuvers and perform equipment

checks. COPDGene® will monitor and provide periodic feedback on each technician’s

performance.

c.Equipment: Leaks in the system, differences in temperature, and poor calibration are all

factors that affect the quality of the test results. Daily spirometer calibration checks should be

performed using a 3.00 Liter syringe as the “gold standard.” Refer to section 14 of the

EasyGuide™ users’ manual for instructions.

d.Analysis: A combination of all factors may affect the quality of the results. Results from

the calibration factors, technician’s impression of test session quality, and the QC supervisor’s

COPDGene®

Version5.0 June 15, 2009

– 49 –

Manual of Operations

Spirometry

impression of test session quality are all integrated to obtain the final FEV1 and FVC results used

in the study.

9n.

Spirometry Calibration, Maintenance, and Hygiene

The EasyOne™ spirometer is designed to minimize the need for cleaning and maintenance (see

sections 13 and 14 in the EasyGuide™ Users’ Manual). The surface of the spirometer and cradle

may be cleaned by wiping with a damp cloth. If a more thorough cleaning is desired, the

spirometer and its Spirette™ cavity may be cleaned with an alcohol wipe or a soft cloth that has

been lightly moistened with isopropyl alcohol. Do not let liquids flow into the Spirette™

cavity of the spirometer while cleaning. The disposable Spirette™ eliminates the need for

cleaning the spirometer between patients. The Spirettes™ are designed for single patient use

only, and must be removed and disposed of after each patient.

Additional guidelines for hygiene and infection control are provided by the American Thoracic

Society and include the recommendation that the technician and patient wash their hands after

testing and that proper attention be given to environmental controls in settings where

tuberculosis or other diseases spread by droplet nuclei are likely to be encountered. Participants

with evidence of obvious upper respiratory infections should not be tested, but rather rescheduled

for testing at a later date.

Beyond battery replacement and the calibration check described below, no periodic maintenance

is required or recommended on the spirometer or cradle. No service should be performed on the

spirometer except by manufacturer-authorized personnel.

On each day a subject will be test, and prior to the first participant tested, the spirometer

calibration should be checked with a 3.00-Liter calibration syringe that has been stored next to

the spirometer. The calibration procedure requires the ndd calibration adapter to connect the

syringe with the spirometer.

To clean the nose clips, they should be wiped with alcohol after each use or discarded.

9o.

Trouble Shooting

Refer to the section 15 of the EasyGuide™ Users’ Manual for troubleshooting tips.

9p.

ndd EasyOne™ Configuration Settings for Use in COPDGene®

When setting up your EasyOne™ spirometer, use the following settings for compliance with

COPDGene® protocol.

Test Settings:

1) Predicted = “NHANES III”

2) Add Ped= “Blank”

3) Value Sel = “Best Value”

4) Interpretation = “- - - -”

5) Lung Age = “OFF”

6) Automatic QC = “ON”

COPDGene®

Version5.0 June 15, 2009

Manual of Operations

Spirometry

– 50 –

7) FVC Selection = “FVC”

8) PEF unit= “L/sec”

9) African ethnic = 88%

10) Asian = 100%

11) Hispanic= 100%

12) Other ethnic = 100%

13) Curve Storage = “3 Best Curves” <Very Important!!! >

General Settings:

1) Time Form = 24 hr(

Date Format = “DD/MM/YY”

2) Date = Enter correct date

3) Time = Enter correct time

4) Alpha-ID = “YES”

5) Tech ID= “YES”

6) Syringe Vol = 3.0 Liter

7) Height Unit = “m/cm”

8) Weight Unit = “kg”

9) Age/Birth = “AGE”

10) Contrast= 8 or adjust as needed

11) Language = “ENGLISH”

12) Altitude= 0 or set to approximate (100 meter increments)

13) Op Mode = “DIAGNOSTIC”

14) Temp = “C”

15) Rel Humidity = Best average guess (0 to 100%)

Report Settings:

1) Printer Type = ”Via PC”

2) Result Data = “3 Best Data”

3) Number of Curves = “3 Best Curves”

4) Graph Types = “Small FV & VT”

5) Header 1 to 4 = “Text that you want to print on the report” Some centers may choose

to enter their Center’s 3 letter code and COPDGene

9q.

Spirometry Checklist

Clinical Center Coordinators are encouraged to print and use this checklist as an aid during

performance of the spirometry testing.

COPDGene®

Version5.0 June 15, 2009

– 51 –

Manual of Operations

Spirometry

Subject Identification Number:

Action

Calibrate spirometer

Safety Assessment Form administered

Prepare to conduct spirometry

Confirm race/ethnicity of participant

Explain purpose of test

Demonstrate maneuver

Complete pre-bronchodilator test

Administer bronchodilator

Wait 15-40

Complete post-bronchodilator test

Print results

□

□

□

□

□

□

□

□

□

□

□

Test Date:

Comments

9r.

Common Questions and Answers for the ndd Spirometer

This document summarizes various problems that have been reported by sites when using the

ndd spirometer. If you are experiencing problems with the ndd, review this document to see if

your problem is described and see if the proposed solution takes care of the problem. If not, then

contact the PFT Core or ndd directly.

1. CALIBRATION CHECK ERRORS

Problem: Unable to obtain an acceptable calibration check.

Explanation: It is not possible for the user to change the calibration in the EasyOne™.

EasyOne™ will not change its calibration since it uses ultrasound transit time

measurement. To enable customers to check the calibration, however, EasyOne™ includes

a calibration check program. The procedure for this cal check is described in the manual.

Potential reasons for failed calibration checks are summarized below, with potential

solutions noted below.

a. The calibration syringe may be defective.

Solution: The syringe itself may be defective, or it may require a special adaptor to

properly attach to the Spirette™. Hans Rudolph makes a calibration syringe specifically

designed for use with the ndd, and we recommend that sites use this syringe. An advantage

of the Hans Rudolph syringe is that it can be properly connected to the Spirette™ without

the need for a flexible tube, which can be an additional source of calibration errors.

b. Faulty ndd device.

Solution: We are aware of at least one instance where the ultrasound transducer was not

mounted correctly in the factory and therefore resulted in excessive variability (up to 10%)

COPDGene®

Version5.0 June 15, 2009

– 52 –

Manual of Operations

Spirometry

during calibration tests. If you suspect you may have a faulty device, notify the Operations

Center, which will in turn notify ndd. If the unit is indeed defective, ndd will replace it.

c. The grey calibration adapter is broken or has been used without the necessary

second screen inside. This can cause an increased variation in the calibration

check.

Solution: This problem only applies for instruments with a serial number below 46000.

All EasyOne™ with serial number 46000 and higher use an improved calibration adapter

design that solves this problem. However if you have an older device this may still be an

issue and we recommend that you use a more recent instrument if possible.

d. The room is very ‘windy’ because the windows and the doors are open to supply

cooling.

Solution: This can interfere with the calibration checks, especially regarding the

inspiratory values. During the cal check staff should make sure that there is no noticeable

air draft in the room.

e. Additional reasons for an inadequate calibration check are that the baseline was

not set correctly at the beginning of the cal check (see below) or that the syringe

was not pulled in or out completely

Solution: Check for these issues and repeat the calibration.

In addition to the calibration check routine, it is also possible to do a calibration check in

the FVC program. However, this is discouraged for the following reasons:

There is a 4% BTPS correction built into the test that causes an over estimation of the

flow and the volume by 4%. The reason for this correction is that the gas in the lung is

warmer than when it arrives in the sensor. This cooling down of the gas also causes a

volume reduction.

FEV1 values will be different since in human tests the FEV1 calculation needs to follow

ATS algorithms. If there is slow flow before the actual start of the blow, this would

cause a lower FEV1. To correct for this, the t-zero point is calculated by back

extrapolation from the maximum volume increase in the Volume Time Curve. FEV1

values are therefore hardly comparable between the syringe and the EasyOne™. The

better the syringe stroke simulates a human spirometry test, the closer the two FEV1

values will be.

The flow parameters (PEF) can usually vary more than the Volume parameters. In the

ATS waveform test a variation of +/- 12% is allowed for PEF. This is because gas can

be compressed and resistance between syringe and sensor can influence the result.

2. SETTING ZERO FLOW BASELINE

Problem: The EasyOne™ does not stop at the end of the test, or volume parameters are

inaccurate in the calibration.

COPDGene®

Version5.0 June 15, 2009

– 53 –

Manual of Operations

Spirometry

Explanation: Prior to every FVC test or calibration check, the zero flow baseline needs to

be set. You should see the message “setting baseline. Block Spirette™ to avoid flow” on

the screen. The technician should make sure that there is absolutely no flow through the

Spirette™ while the baseline is set. Hold the instrument still and block the Spirette™ or

hold the instrument close to the body to make sure there is no flow. If the baseline is not

set correctly, there is an offset in the flow measurement. This causes an inaccuracy in the

FVC value. The longer the measurement takes, the bigger the error in the result. If there is

a 20 ml offset, the error after 8 seconds in the FVC value is 160 ml. In the FEV1 value the

error is only 20 ml.

If EasyOne™ does not stop at the end of the maneuver, there is an offset in the baseline

and the instrument continues to accumulate volume. If the volume-time curve shows a very

long test, and there was a very low volume accumulation over many seconds, the baseline

was probably not set correctly.

3. SMALL CURVES

Problem: Technicians have sometimes observed that there are sometimes small curves

between normal maneuvers. Quality Grades were poor even though the maneuvers look

good.

Explanation: The test was started too late. The instrument was started by pressing ENTER

after the patient had already started blowing. Only a part of the maneuver was captured by

the EasyOne™. This could happen because the instrument was not held by the patient but

by the technician. Most of these curves are not selected as best curves since the FVC is too

low. The flow volume curves look strange on the printout since they start immediately with

the peak flow, and of course it was smaller since the total volume measured was smaller.

This problem can be very frustrating for the participant since it triggers a bad

reproducibility of maneuvers message even though the patient gave his best.

This problem can be easily avoided if the patient holds the instrument himself. The

instrument is very robust and there is no need to fear that the patient might break it, for

instance by dropping it. The technician should start the instrument and hand it over to the

patient. EasyOne™ recognizes when the actual test starts. EasyOne™ will recognize the

incorrect trial and give a coaching message.

4. “ERROR 6” MESSAGE

Problem: Some instruments have occasionally displayed an “error 6” message on the

screen. The message prevents further test maneuvers from being performed on the patient.

Explanation: This problem should have been fixed with the Firmware Version 1.20. The

“error 6” message means that a database conflict came up without reason. This should only

be possible in very rare cases, but ndd reports that it could reproduce the effect.

Fortunately no data are lost when the message comes up, although as noted it is not

possible to add further trials or a post test to the session. If the problem does occur, it is

strictly a software issue and does not indicate a problem with the spirometer device itself.

COPDGene®

Version5.0 June 15, 2009

– 54 –

Manual of Operations

Spirometry

5. EASYONE™ WON’T ACCEPT ADDITIONAL TRIALS

Problem: Some technicians have reported that they have tried to add more trials but the

equipment does not seem to take the new data.

Explanation: EasyOne™ allows up to 15 trials per test session. Since ATS recommends

stopping after 8 trials if no good results can be achieved, 15 seems to be a safe number. If

after 5 trials, the patient has reached a QC Grade C, the instrument considers the test

session to be sufficient and suggests ending the test. However, it is certainly possible to

add more trials at this point.

One reason for this perceived problem may be a misunderstanding of the device display

and its operation: The following screen is shown after a trial:

The information on the display depends on the test status:

If the test is not finished the display shows the best and the current curve. The

numeric results to the right are also from these curves.

If the test session is finished, the results from the two best curves are shown.

This means that if a curve is added after the EasyOne™ thinks the session is finished and

the result of this trial is lower than the two previous best results, the screen only shows the

two best tests. Staff also need to understand the criteria EasyOne™ uses to determine a test

session is finished. The ndd software considers the session to be finished, and hence

recommends stopping:

As soon as a quality grade of A or B is reached, or

If, after 5 trials, a quality grade of at least C has been reached

As noted previously, even though the ndd deems the session to be finished, the technician

can still add additional tests, up to a maximum of 15, in order to get a better QC grade. The

technician just needs to realize that the results from the new curve will only be displayed in

this case if it is one of the two best curves in the session. Even if the display does not

change, the curve is stored in the database. Only when 15 trials are reached it is not

possible to add any more trials.

As soon as a POST test has been started, it is no longer possible to add trials to the PRE

test. However an additional 15 trials may be conducted for the POST test, for a total of 30

overall. (see also #10)

6. “KEEP GOING...” MESSAGE

Problem: Sometimes the tests do not stop and the “Keep going…” message is shown.

Explanation: In most cases this happens due to incorrect baseline setting (see #2 above).

At the beginning of a new test session make sure the user blocks the Spirette™ to avoid

COPDGene®

Version5.0 June 15, 2009

Manual of Operations

Spirometry

– 55 –

flow. If this is not done carefully, a wrong flow offset can be set which leads to the

described behavior.

7. POWER DOES NOT TURN ON/OFF CORRECTLY

Problem: the device cannot be switched on or off.

Explanation: ndd has not been able to reproduce this problem, and would appreciate more

information in order to properly investigate the issue. If the problem occurs, please supply

the following information: does the issue always happen with the same device, how often

does it happen, and under what circumstances does it happen? In the meantime, if the

problem does occur it should be remedied by replacing the battery. No data can be lost by

removing the batteries.

8. FLOW-VOLUME CURVES HAVE STRANGE SHAPE

Problem: Some of the flow-volume curves have a strange shape and the FEV1 and FVC

values are low (see curve marked with red and heavy line below). (See also #3 above)

Explanation: This happens if the patient blows before the device is ready to accept flow; it

should not happen for the very first trial since the baseline is required to be set for this

maneuver. The problem appears to result from the manner in which the tests are performed

at some sites. The interviewer would hold the device while the patient inhaled through it

prior to the exhalation. At the end of the inhalation the interviewer would press the

ENTER key to start a new trial.

The ndd has a delay of approximately one second between pressing the ENTER key and

the device being ready to measure flow. The spirometer beeps when it is ready. If the

patient is already exhaling when the key is pressed or shortly after that, this leads to the

unusual curves. This also leads to strong variations in the measurement results.

These bad curves are usually not accepted in the final results since the parameters are all

low. However it is frustrating for the interviewer and the patient when they do not get

acceptable results. The interviewer should make sure that the instrument is started early

enough, at least one second prior to the forced exhalation. (see also #3 above)

COPDGene®

Version5.0 June 15, 2009

– 56 –

Manual of Operations

Spirometry

9. QUALITY GRADE GETS WORSE

Problem: It can happen that the quality grade changes from A or B to C (i.e., gets

‘worse’).

Explanation: This can happen if an acceptable trial is added that has substantially better

values than the previous best trial, thus causing the variability between the two best

maneuvers (the difference between the two best values) to increase. Depending on the

magnitude of the difference, the quality grade can even change to a C or D. (see also #5

and #11)

In some cases the screen will display the message “Deeper Breath” when this happens. In

this particular case the message is misleading and should probably state, “Great effort. Do

another just like it!” The technician should be aware of this and coach the participant

accordingly. ndd hopes to change this message in a future version of the software.

10. EASYONE™ WON’T ACCEPT POST TEST DATA

Problem: Sometimes the wait after administering a BD is so long that the device

automatically switches off. If the test is recalled then, it is sometimes impossible to do a

POST test and only the selection NEW and EXIT is available instead of ADD, POST and

NEW.

Explanation: This effect is caused not by the time delay, but by the way the software

currently handles PRE tests with 15 trials. Due to a software error, no POST test can be

added when doing a recall of a pre test with 15 trials. This problem does not occur as long

as there are less than 15 trials in the pre session. This problem has been fixed as of

Firmware Version 1.20 of the EasyOne™ software.

Please alert the Operations Center if this occurs with the current version of EasyOne™ or

if it occurs in other circumstances than described above with earlier versions of the

software.

11. EASYONE™ ACCEPTS BAD TESTS

Problem: Occasionally the EasyOne™ accepts bad curves, which can result either in

inappropriate values for FEV1 and FVC or in inappropriate QC grades.

Specific examples of the situations in which apparently unacceptable curves have been

reported as being accepted by the EasyOne™ software are summarized below.

a. A double inspiration due to cough or hesitation in starting

b. Variable effort or cough during a trial

Explanation: ndd is aware of this problem and is working on a fix. Please notify the

Operations Center of all such instances as they occur so that we can pass them onto ndd for

further analysis. In versions of the EasyOne™ software prior to 2.3, there is no way to

allow technicians to exclude inappropriate tests on their own. In version 2.3, it is possible

to switch off the QC grade function under the “Configuration” menu. When you do this,

the device allows you after every trial to reject an acceptable trial or to accept an

unacceptable trial. The disadvantage of this option is that you don’t see the QC grades

COPDGene®

Version5.0 June 15, 2009

– 57 –

Manual of Operations

Spirometry

during the test and on the report. Other than this, the test procedures and messages are the

same as when the QC grade function is turned on.

With the EasyWare™ you are able to see the QC grades on the screen. (Note that any

decisions regarding acceptability or unacceptability may be overridden by the spirometry

reading center.)

12. STAFF DISAGREE WITH GOOD QC RATING

Problem: Even though tests display a grade of A, the technician may disagree.

Explanation: ndd is aware of this problem and is working on a fix. Please notify the

Pulmonary Function Core of all such instances as they occur so that we can pass them onto

ndd for further analysis (see also #11). In the meantime it is likely that such curves will be

detected as part of the central QC review.

Related Pulmonary Function Documents:

1.American Association for Respiratory Care Clinical Practice Guideline: Static lung

volumes. Respir Care 1994;41:629-636.

2.Clausen JL, Coates A, Quanjer PH. Measurement of lung volumes in humans: reviews and

recommendations from an ATS/ERS workshop. Eur Respir J 1997;10:1205-1206.

3.Coates AL, Reslin R, Rodenstein D, Stocks J. Measurement of lung volumes by

plethysmography. Eur Respir J 1997;10:1415-1427.

4.Newth CJ, Enright P, Johnson RL Jr. Multiple-breath nitrogen washout techniques:

including measurements with patients on ventilators. Eur Respir J 1997;10:2174-2185.

5.American Thoracic Society: Lung function testing: selection of reference values and

interpretive strategies. Am Rev Respir Dis 1991;144:1202-1218.

6.Stoller JK, Basheda S, Laskowski D, Goormastic M, McCarthy K. Trial of standard versus

modified expiration to achieve end-of-test spirometry criteria. Am Rev Respir Dis

1993;148:275-280.

7.Enright P. Can we relax during spirometry? Am Rev Respir Dis 1993;148:274.

COPDGene®

Version5.0 June 15, 2009

– 58 –

Manual of Operations

Spirometry

10.

Six-Minute Walk Test

The six-minute walk test is a timed walk involving a familiar activity and requiring minimal

technical resources. It has been shown to be a reproducible objective indicator of functional

capacity and is a part of the multi-component BODE index that correlates with survival. It is

important to emphasize that this is a test of maximum exercise performance; subjects should be

instructed and encouraged to push themselves to achieve maximal distance. Only one walk test

will be done on each subject.

10a. Contraindications

•

•

•

•

Proven or suspected unstable coronary artery disease or angina

Exercise related syncope

Proven or suspected claudication

Uncontrolled hypertension (resting systolic blood pressure > 200 mmHg or resting

diastolic blood pressure > 110 mmHg)

• Resting bradycardia (< 50 beats/min), history of complex ventricular arrhythmia or

sustained SVT

• Use of > 6 L/min oxygen flow with activity

10b. Facilities/Equipment

•

•

•

•

Stopwatch

Course pre-marked with distances

Portable oxygen delivery system if the patient is on oxygen (nasal cannula)

Subject instructions below

10c. Walk Course

The path should be unobstructed, flat, and indoors. If a public corridor is used, ability to control

traffic should be assured. Each site is encouraged to use the same course for all subjects. The

course description should be recorded on the walk test data collection form. The course should

be designated as straight or circular, and the distance of the course should be included on the

walk test form.

10d. Procedure

Oxygen saturation, heart rate, and blood pressure will NOT be monitored during the walk.

A staff member will carry a portable oxygen supply if supplemental oxygen is used by the

patient.

10e. Patient Preparation

• Short acting bronchodilator may be used if so prescribed as part of the subject’s usual

medications, but if taken they must be used at least 20 minutes before testing.

• A light meal 2 to 4 hours prior to testing is advised; at least 2 hours must have elapsed

since the patient last ate a meal.

COPDGene®

Version5.0 June 15, 2009

– 59 –

Manual of Operations

Six-Minute Walk Test

• The patient should rest in a sitting position for 10 minutes before testing.

• Patient should wear loosely fitting clothing and comfortable walking shoes.

10f. Subject Instructions

Pre-test instructions: Identical instructions should be given to each subject. Say to the subject:

“You are now going to begin a six minute walking test. The object of this test is to walk as

quickly as you can to cover as much ground as possible in six minutes. You may slow down if

necessary. If you stop, we wish you to continue the walk again as soon as possible. Your goal is

to walk as fast and as far as you can in six minutes.” Then review the course with the patient.

Specific instructions on walking the course will differ somewhat between sites because of

differences in the layout of each walking course. Finally, say “Wait until I say ‘Start’ before

beginning. It is important that you not talk during the test unless you are having a problem. Do

you have any questions?”

10g. Instructions during Walking

At the end of each minute the patient will be given the time elapsed, the time remaining, and a

standard phrase of encouragement as follows:

Minute 1 – “Do your best.”

Minute 2 – “Try your hardest.”

Minute 3 – “Keep going.”

Minute 4 – “Give it your all.”

Minute 5 – “Walk faster if you can.”

For example, at the end of minute 4, the technician would say: “4 minutes, you have 2 to go, give

it your all.”

10h. Research Coordinator Responsibilities

Carry supplemental oxygen, if oxygen is routinely used by the subject during activity. Begin by

saying “Start.” Provide standard instructions before the test and standard information and

encouragement during the walk at set intervals as described above. Say “Stop” at the end of six

minutes. Measure the distance walked to the nearest meter or foot. Rest periods are not recorded;

a test lasts 6 minutes if the patient is on the course for 6 minutes; the patient need not walk

continuously for 6 minutes for the walk to be considered to have a normal termination.

If the test terminates abnormally (i.e., patient leaves the course before 6 minutes have elapsed),

record the distance walked, time on the course, and reason why the test was terminated before 6

minutes had elapsed.

10i. Abnormal Test Termination Criteria

•

•

•

•

Chest pain suspicious for angina.

Evolving mental confusion or lack of coordination.

Evolving lightheadedness.

Otherwise warranted based on clinical condition.

COPDGene®

Version5.0 June 15, 2009

– 60 –

Manual of Operations

Six-Minute Walk Test

11.

Imaging

11a. CT Acquisition at Study Sites

11a.1.

Site Survey

Clinical Centers must complete an initial Imaging Site Survey Form (Appendix) to identify

responsible personnel and CT scanner models prior to obtaining scans for the study.

Each scanner will be assigned a specific ID number by the Imaging Core; this information will

be recorded on the Image Collection Form when each scan is performed.

Survey should be faxed to the Imaging Center at 303-270-2538. This survey should also be kept

on file at each clinical site.

11a.2.

Technologist Training

A PowerPoint training set, “CT Imaging in COPDGene®” has been implemented for

technologists at the Clinical Centers to assure understanding of the outline of the study, and

compliance with the radiology protocols. Each technologist involved in the acquisition of scans

must be certified as having reviewed the training set prior to performing CT scans for the study.

After the technologist reviews the PowerPoint slides, they must sign the Technologist Training

Log to attest that they have reviewed and understand the study protocol. Before the start of

enrollment, the clinical center must ensure that each involved technologist reviews the slides and

signs the Log. After all technologists have reviewed the slides, the clinical center should fax this

form to the Imaging Core at 303-270-2538. The original copy of this form should be kept at the

clinical sites.

When additional radiologic technologists become involved in the study, the additional

technologist should review the slides and sign the log. At this point, the clinical center should

again fax the form to the Imaging Core so that the core has record of every new technologist.

The Imaging Core will keep a record of the study-affiliated technologists at every clinical center.

An email will be sent from the Imaging Core to the study coordinator notifying them regarding

the receipt of a signed Technologist Training Log.

11a.3.

Scanner Quality Assurance

At each study site, each CT scanner used in the study will scan a phantom at the initiation of

subject enrollment, on an ongoing monthly basis, and after any hardware or software change.

The site will use a customized COPDGene® phantom.

The phantom will be scanned using the study protocol parameters for each scanner. The phantom

scans will be transferred on DVD to the Imaging Core for analysis. Scan information will be

recorded on the Image Collection Form (see appendix).

COPDGene®

Version5.0 June 15, 2009

– 61 –

Manual of Operations

Imaging

The DVD with the Phantom Scan information and the Image Collection Form should be sent to

the Iowa Imaging Core.

11a.4.

1.

2.

3.

Subject Preparation for CT Scan

4.

Prior to the study, the subject’s identity will be confirmed according to institutional policy.

The subject will remove all metallic devices from the chest area.

The subject will be informed of the importance of compliance with the breathing

instructions (in Section 11c below). Ability to comply with instructions should be assessed,

and conditions that might impair compliance such as deafness, breathlessness, or mental

impairment should be noted.

At least one rehearsal of the end-inspiratory breathhold should be performed.

CT Acquisition

11a.5.

All CT scans will be obtained using the protocol and breathing instructions as indicated (see

Protocol in Section 11c). Scans must be reconstructed using two algorithms, edge enhancing and

smooth. Contiguous end-expiratory CT images will also be obtained where possible. Additional

reconstructions may be performed as required at study site, but only the contiguous thin section

images, reconstructed in two formats, should be sent to the Core Laboratory.

The responsible CT technologist should complete the Image Acquisition Form, and sign to

confirm that the study protocol was followed and the scan meets the expectations of the study.

The protocol for CT acquisition will be printed on the reverse of the Image Collection Form.

11a.6.

Non-study Scans

Scans on eligible patients performed outside the COPDGene® study may be accepted for

purposes of the analysis if they meet the non-study scan criteria (see appendix). Final decisions

on acceptability of these scans will be at the discretion of the Imaging Core, based on their

adequacy for quantitative image analysis. The study coordinator should ensure that the CT

Assessment Scoresheet and the Non- COPDGene® Scan CT Acquisition Parameters forms are

completed for these subjects (see instructions in the CT Assessment Scoresheet section).

11a.7.

CT Assessment Scoresheet

The radiologist will complete the CT Assessment Scoresheet on each subject for both study and

non-study scans, and will sign off on the quality of the scan. This scoresheet will be both

submitted electronically and the original, signed copy will be sent with the Image DVD to the

Imaging Core.

How to Submit the CT Assessment Scoresheet Electronically:

1.

2.

Sign in to the COPDGene® website at

https://biosweb.njc.org/sec/COPDGene/MainPage.cfm

Click on the “Forms” heading.

COPDGene®

Version5.0 June 15, 2009

– 62 –

Manual of Operations

Imaging

3.

4.

5.

6.

7.

Download the scoresheet as a PDF by clicking on the “CT Assessment Scoresheet” link.

Complete the required information on the form.

Print and save a copy of this form at the Clinical Center.

Submit the scoresheet information by clicking on the “Send the data to the DCC” link at

the bottom of the last page.

A copy signed by the radiologist must be shipped to the Imaging Core with the Image

DVD.

Note: The electronic submittal of the CT Assessment Scoresheet is performed by either the

radiologist or the study coordinator. The process will be determined and coordinated individually

by each clinical center.

11a.8.

Image Acquisition Form

There are a number of Image Acquisition Forms provided on the COPDGene® website. The

purpose of the different versions is to provide the technologist with the form and the appropriate

protocol for the scanner that will be used. The proper version of the Image Acquisition Form will

help to ensure that the proper COPDGene® protocol is accurately followed.

The Image Collection Form should be printed from the website by the coordinator and a bar code

applied to the top portion. The form with the appropriate protocol should then be submitted to

the radiologist technologist before the scan is performed. The technologist will fill out this form.

When the Images are shipped to the Imaging Core, this form should be included in the shipment.

If the CT was not performed according to the COPDGene® protocol, the “Non- COPDGene®

Scan CT Acquisition Parameters” form must be completed by doing the following (to be

performed by the coordinator):

1.

2.

3.

4.

5.

6.

Sign in to the COPDGene® website at

https://biosweb.njc.org/sec/COPDGene/MainPage.cfm

Click on the “Forms” heading.

Download the form as a PDF by clicking on the “CT Acquisition Parameters” link.

Complete the required information on the form.

Print and save a copy of this form at the Clinical Center.

Submit the scoresheet information by clicking on the “Send the data to the DCC” link at

the bottom of the last page.

Scan Shipment

11a.9.

CT Scans on DVD

All protected health information (PHI) should be removed from the DICOM header, and

replaced with the participant’s study ID. Specific DICOM fields to be removed include the

patient name, accession number, medical record number, date of birth, date of examination, and

comment field.

COPDGene®

Version5.0 June 15, 2009

Manual of Operations

Imaging

– 63 –

The anonymized scan data should be written to DVD, and sent to the Imaging Core in a

polypropylene protective case.

Ideal Items to Be Used:

Item

DVD-R (Gold Archival Grade)

Polypropylene DVD Case

Padded Pak

Preprinted Airbill

Manufacturer

Verbatim Ultralife Gold Archival

Grade 4.7 GB 8X

TRIMPak II

FedEx

FedEx

Provided By

Clinical Center

Imaging Core

Imaging Core

Imaging Core

Imaging Shipment Form

The Imaging Shipment Form should be included with each shipment sent to the Imaging Core.

This Form will be used to document the materials included in each shipment as a quality control

measure. Be sure to add the barcode of each subject to this form.

Keep a copy of this form at the Clinical Center.

Labeling Imaging Materials

The DVD should be labeled using a black permanent marking pen. The following information

should be printed clearly on the DVD:

• Subject ID including Clinical Center, such as NJC

• Date of scan

• COPDGene®

The DVD should then be placed in the Plastic Case provided. Only one DVD should be placed in

each Plastic Case. A barcode label should be placed on the Plastic Case. Barcodes are generated

by the DCC (these are the same barcodes that are used for Blood shipment). These barcodes will

be emailed to the study coordinator as a PDF file after the subject is assigned an ID number by

the DCC on the COPDGene® website. The PDF file will be printed on Avery 5267 8.5 x 11

paper labels, 4 across and 20 down, using a standard office printer. An example of the barcode is

shown below.

Barcodes should be placed on:

• CT Scan Plastic Case

• CT Scan Assessment Scoresheet

COPDGene®

Version5.0 June 15, 2009

– 64 –

Manual of Operations

Imaging

• CT Acquisition Form or Non-Study Image Acquisition Form

• Image Shipment Form

A backup DVD, labeled in the same manner as the primary DVD sent to the Imaging Center,

should be stored at the clinical site with the subject’s study records and all source documents.

Shipping CT Scans to Imaging Core

Upon initiation of enrollment at the clinical center, the Imaging Core will arrange a shipment of

25 plastic DVD cases, 25 FedEx Padded Paks and 25 preprinted airbills to the site for shipment

of CT scans.

Shipping Instructions:

1.Write information on DVD as noted above.

2.Place DVD in plastic case.

3.Place barcode labels on materials as noted above

4.Complete the Image Shipment Form.

5.Place scan DVD in the case, CT Acquisition Form, Image Shipment Form, and CT

Assessment Scoresheet in FedEx Padded Pak.

6.Complete the preprinted FedEx Airbill and place on sealed package.

7.Call FedEx for a pickup, or drop-off at local Fed-Ex location.

8.Enter the shipment tracking number, shipment date, and subject IDs on the COPDGene®

Website by clicking on the “CT Data” under the Tracking heading. This must be done on

the same day that the CT is shipped.

9.The first 10 CT scans from each site or subsite should be sent as soon as they are

completed. After the first ten have been quality checked, subsequent scans can be sent in

weekly batches with a maximum of ten DVDs in each package.

10. More shipping supplies will be sent to your site as needed. Contact the Imaging Core at

303-270-2529 when your supply runs low.

11a.10. Quality Assurance

Anonymized images will be submitted on DVDs to the Imaging Core in DICOM format, using a

study ID as the only identifier. Upon arrival at the Imaging Core at National Jewish Medical and

Research Center, CT media will be processed by the research staff to verify anonymization and

appropriate identification of study information, protocol compliance, image quality, and image

count.

The Imaging Core will verify that the forms match the ID on the DICOM header on the

submitted DVD. If quality issues are identified, the staff will complete a quality form and will

contact the site to attempt resolution. The site will be notified within 5 business days of quality

problems. If the CT data are found to be unusable because of quality problems, the participant

will not be enrolled in the study.

11b. Data Reporting

The following data will be reported to the DCC:

• Receipt of complete and technically adequate scan

COPDGene®

Version5.0 June 15, 2009

Manual of Operations

Imaging

– 65 –

• Patient dose and scan duration as recorded on the image acquisition form

• Quantitative parameters outlined above

• Quality assurance data

11c. Instructions for CT Scan Acquisition

GENERAL: This study consists of 2 scouts (topograms) and 2 scans. All scans use the same

parameter grid.

CONTRAST:

Oral/IV. None.

SUPINE INSPIRATION: Start at bottom of lungs, end at top of lungs. Instruct the patient

to breathe as follows:

“For the first part of this study you will be asked to hold your breath in for about 20

seconds. If you cannot hold your breath that long, try the best you can and then

take very shallow, slow breaths if you need to.”

“For now, take several easy, deep breaths and relax while we prepare to take a CT

scan of your lungs.”

Allow patient to breathe and relax for at least 15 seconds.

“I am now going to give you specific breathing instructions. Try to follow as best

you can.”

“Take in a deep breath….and let it out.”

“Take in another deep breath….and let it out.”

“Take in another deep breath, and hold your breath in. Keep holding your breath!”

Scan the patient in one breath-hold at full-inspiration.

When the scan is completed, tell the study participant to “Breathe and relax!”

SUPINE EXPIRATION: Same protocol as SUPINE INSPIRATION. Start at bottom of

lungs, end at top of lungs. Instruct the patient to breathe as follows:

“For the second part of this study you will be asked to blow out your breath and hold

it out for about 20 seconds. This is usually more difficult than holding your breath

in, but do the best that you can. If you cannot hold your breath out that long, take

a very slow shallow breath in if you need to.”

“For now, take several easy, deep breaths and relax while we prepare to take the last

CT scan of your lungs.”

Allow patient to breathe and relax for at least 15 seconds.

“I am now going to give you more specific breathing instructions. Try to follow as

best you can. ”

“Take in a deep breath….and let it out.”

“Take in another deep breath….and let it out.”

“Take in another deep breath, let it out and hold your breath out! Do not breathe!”

Scan the patient in one breath-hold at expiration as quickly as possible.

When the scan is completed, tell the study participant to “Breathe and relax!”

COPDGene®

Version5.0 June 15, 2009

– 66 –

Manual of Operations

Imaging

(b) Inspiratory CT

GE

LS 16

Helical

See mA

16x0.625

1.375

13.75

120

400 @ 0.5s

Effective mAs:200

CAREDose 4D off

CAREDose 4D off

Off

Effective mAs:200

400 @ 0.5s

mAs 200

120

120

120

120

13.75

13.2

21.1

0.5

1.375 mm

1.1

1.1

1.188

0.923

0.5

120

mAs 200

64x0.625

16x0.75

64x0.6

16x0.75

40x0.625

See mA

0.5

0.5

0.5

0.5

VCT Helical

Spiral

Spiral

Axial Helix

Axial Helix

VCT-64

Sensation-16

Sensation-64

16 slice

40 slice

GE

SIEMENS

SIEMENS

PHILIPS

PHILIPS

PHILIPS

64 slice

Axial Helix

0.5

64x0.625

0.923

0.5

120

mAs 200

Scanner make

Scanner model

Scan Type

COPDGene®

Version5.0 June 15, 2009

Auto-mA off

Off

Off

Off

BONE

B46f

0.75

0.5

Lungs*

0.625

0.625

Lungs*

Lungs*

0.625

0.625

BONE

B46f

0.75

0.5

Lungs*

Detail (D)

0.9

0.45

Lungs*

Detail (D)

0.9

0.45

Lungs*

Detail (D)

0.9

0.45

Lungs*

Standard

0.625

0.625

Lungs*

0.625

0.625

Lungs*

Standard

B31f

0.75

0.5

Lungs*

B31f

0.75

0.5

Lungs*

B

0.9

0.45

Lungs*

B

0.9

0.45

Lungs*

B

0.9

0.45

Lungs*

Rotation Time (s)

Det. Configuration

Pitch

Speed (mm/rot)

kVp

mA

– 67 –

Dose modulation

Reconstructions

RECON1

Algorithm

Thickness (mm)

Interval (mm)

DFOV (cm)

RECON 2

Algorithm

Thickness (mm)

Interval (mm)

Manual of Operations

Imaging

DFOV (cm)

* reconstruction field of view should encompass the widest diameter of the lung.

(c) Expiratory CT

GE

LS 16

Helical

See mA

16x0.625

1.375

13.75

120

100 @ 0.5s

Effective mAs:50

CAREDose 4D off

CAREDose 4D off

Off

Effective mAs:50

100 @ 0.5s

50 mAs

120

120

120

120

13.75

13.2

21.1

0.5

1.375 mm

1.1

1.1

1.188

0.923

0.5

120

50 mAs

64x0.625

16x0.75

64x0.6

16x0.75

40x0.625

See mA

0.5

0.5

0.5

0.5

VCT Helical

Spiral

Spiral

Axial Helix

Axial Helix

VCT-64

Sensation-16

Sensation-64

16 slice

40 slice

GE

SIEMENS

SIEMENS

PHILIPS

PHILIPS

PHILIPS

64 slice

Axial Helix

0.5

64x0.625

0.923

0.5

120

50 mAs

Scanner make

Scanner model

Scan Type

COPDGene®

Version5.0 June 15, 2009

Auto-mA off

Off

Off

Off

BONE

B46f

0.75

0.5

Lungs*

0.625

0.625

Lungs*

Lungs*

0.625

0.625

BONE

B46f

0.75

0.5

Lungs*

Detail (D)

0.9

0.45

Lungs*

Detail (D)

0.9

0.45

Lungs*

Detail (D)

0.9

0.45

Lungs*

Standard

0.625

0.625

Lungs*

0.625

0.625

Lungs*

Standard

B31f

0.625

0.5

Lungs*

B31f

0.75

0.5

Lungs*

B

0.9

0.45

Lungs*

B

0.9

0.45

Lungs*

B

0.9

0.45

Lungs*

Rotation Time (s)

Det. Configuration

Pitch

Speed (mm/rot)

kVp

MA

– 68 –

Dose modulation

Reconstructions

RECON1

Algorithm

Thickness (mm)

Interval (mm)

DFOV (cm)

RECON 2

Algorithm

Thickness (mm)

Interval (mm)

Manual of Operations

Imaging

DFOV (cm)

* reconstruction field of view should encompass the widest diameter of the lung.

12.

Collection of Blood Samples

12a. Purpose of COPDGene® Study Center LAB Manual

This manual serves to identify those involved with the COPDGene® LAB and to clarify standard

operating procedures for work on this study. These procedures should be adhered to by all staff

members and monitored by quality assurance and quality control activities.

12b. Biological Specimen Handling

12b.1.

Labeling and Identification

Purpose: Ensure all collected samples are labeled accurately and adequately.

By Whom: COPDGene® Site Coordinator/phlebotomist

Procedure:

The COPDGene® site will use the COPDGene® information system, as provided by the Data

Coordinating Center (DCC), to print bar code labels that will be used to label all specimen

collection containers, documents, etc. Example of the label designed by the DCC is below.

The labels will be generated at the clinical site using the system developed by the DCC. These

labels will be printed on Avery 5267 8.5 x 11 paper labels, 4 across and 20 down, using a

standard office printer. The items labeled with this label will not require a cryogenic, thermal

transfer label to be applied at the Clinical Center. It will be printed and applied at the JHBR Lab.

These labels will be used for all blood tubes and specimen transmittal documents for the lab.

They can be used also for other documentation.

Place label on tube from top to bottom.

The bar code labels should be placed on both the blood tubes and on the “Blood Specimen

Transmittal Form.”

12b.2.

Blood Draws/Phlebotomy

Purpose: To ensure that venipunctures are performed following standard safety guidelines, and

to ensure that blood samples are collected in the correct order and in accordance with study

protocols.

By Whom: COPDGene® Site/Phlebotomist/nurse

COPDGene®

Version5.0 June 15, 2009

– 69 –

Manual of Operations

Collection of Blood Samples

Procedure:

The following table summarizes draw order and volume for the collection of blood samples into

plastic vacutainer tubes.

Draw Order

1

2

3

4

Tube Type

EDTA purple top

EDTA purple top

EDTA purple top

SST red-grey top

Total

Tube Amount

10 ml

10 ml

10 ml

10 ml

40 ml

Technique:

Blood should be drawn using standard phlebotomy technique with tourniquet applied to identify

vein and released after adequate blood flow is obtained.

Needle or catheter should preferably be no smaller than 21 gauge to minimize risk of hemolysis.

12b.3.

Sample Processing

Immediately after collecting the blood samples, ensure the barcode labels are on each tube and

correspond to the correct study subject

Samples should be refrigerated as soon as possible after they are drawn; samples MUST be

refrigerated within 4 hours. All samples should be placed in a sample storage refrigerator. The

refrigerator should be kept at +2 to +8 degrees C.

Cold Packs

Cold packs must be frozen at -20 degrees C for at least 12 hours prior to use in the shipping

container. Cold packs should be placed flat on a shelf in the freezer.

All Tubes

Invert tubes to ensure specimen integrity.

Place tube in refrigerator after samples have been processed but

prior to shipping to JHBR lab.

Purple Top Tubes

Purple top tube with EDTA should be inverted (not shaken) three

times to mix EDTA Powder thoroughly.

COPDGene®

Version5.0 June 15, 2009

Manual of Operations

Collection of Blood Samples

– 70 –

SST Red-Grey Top Tube

Procedure*:

1) After obtaining the SST sample, allow sample to clot 30 minutes in a vertical position.

2) Follow manual instructions for use of local centrifuge, insuring balance of tubes within the

centrifuge

3) Centrifuge it at 2500 RPM or 1000 to 1300 g for 15 minutes.

4) Do not create aliquots at the clinical center; ship the centrifuged SST sample to the

COPDGene® Biorepository along with the EDTA tubes.

12c. Sample Distribution/Shipping

12c.1.

Shipper Assembling and Packaging Instructions

Purpose: To ensure samples are correctly packaged and shipped according to specimen needs,

study protocol, and IATA guidelines.

General Requirements

Technicians must be familiar with IATA guidelines outlined in the appendix. The packaging

must be of good quality and strong enough to withstand the shocks and loadings normally

encountered during transport – including trans-shipment between transport units and between

transport units and warehouses as well as any removal from a pallet or over pack for subsequent

manual or mechanical handling. Packaging must be constructed and closed so as to prevent any

loss of contents that might be caused under normal conditions of transport, by vibration, or by

changes in temperature, humidity or pressure.

The packaging must consist of three components:

a)A primary receptacle; (blood tubes – EDTA and SST)

b)A secondary packaging; (Plastic biohazard bag and Tyvek® bag – STP-710 along with the

absorbent material – STP-152)

c)A rigid outer packaging. (Styrofoam and the cardboard box – STP-309)

Primary receptacles (blood tubes) must be packed in secondary packaging (plastic and the

Tyvek® bag) in such a way that, under normal conditions of transport, they cannot break, be

punctured or leak their contents into the secondary packaging. Secondary packaging must be

secured in outer packaging (Styrofoam and cardboard box) with suitable cushioning material.

Any leakage of the contents must not compromise the integrity of the cushioning material or of

the outer packaging.

Step by step procedure:

A Power Point presentation outlining specific packing information for the blood tube shipment

has been developed and is available on the COPDGene® web site. All coordinators are expected

COPDGene®

Version5.0 June 15, 2009

Manual of Operations

Collection of Blood Samples

– 71 –

to be familiar with this presentation. Shipping containers must be prepared in accordance with

the steps shown in that presentation.

1) The Styrofoam container should be inserted inside the STP-309 cardboard box.

2) Place a frozen cold pack inside the Styrofoam container.

3) The blood tubes should be inserted into the bubble wrap sleeves and then rolled into a

bundle. Secure the bundle with a rubber band.

4) Then place the bundle in the biohazard plastic bag along with the absorbent material. The

absorbent material does not need to be wrapped around the bubble wrap sleeves.

5) Lay the plastic bag on a flat surface and press hard from the center working outward and

gently seal the bag in order to remove as much air as possible.

6) Place the plastic bag into the Tyvek® bag and seal it.

7) Place the Tyvek® bag flat on the cold pack inside the Styrofoam box.

8) Place the second frozen cold pack on top of the Tyvek® bag. This will limit the space for

movement of the secondary containers within the outer packaging.

9) Place the top on the Styrofoam container.

10) Place the appropriate, completed Blood Tracking Form on the top of the outside of the

Styrofoam container.

11) Finally, securely seal the STP-309 box with packing tape.

12c.2.

Shipping from COPDGene® Sites to JHBR LAB

Purpose: Ensure samples are correctly packaged and shipped according to specimen needs,

study protocol, and IATA guidelines.

By Whom: COPDGene® Site/Study Coordinators

General Procedure:

Note: Samples should be shipped as soon as possible. For the most efficient use of resources,

samples from three subjects should be placed in one shipping container whenever possible.

Samples should not be stored at the clinical site for more than four days prior to shipping;

to meet this requirement, samples from only a single subject may need to be packed in a single

shipping container. Specimens may only be shipped Monday through Thursday to ensure arrival

at the JHU LAB on a weekday. Avoid sending shipments that will be expected to arrive at JHU

LAB on holidays.

COPDGene®

Version5.0 June 15, 2009

– 72 –

Manual of Operations

Collection of Blood Samples

SHIP TO:

Stacey Meyerer

Johns Hopkins University

Bloomberg School of Public Health

615 North Wolfe Street

Room W6618

Baltimore, MD 21205

410-955-7203

1) Refer to the previous section for detailed instructions on assembling and packaging the

shippers.

2) Place the specimen transmittal forms for each subject on the top of foam on the inside of

the shipping box and seal the box.

3) Complete the preprinted FedEx airway bill with DATE, PRIORITY OVERNIGHT,

OTHER PACKAGING and list the number of boxes you will be shipping.

4) Place the airway bill on the box, inside the clear plastic sleeve and call for FedEx pick up.

5) Enter the shipment tracking number on the COPDGene® website, which will generate a

shipment notification that will be sent to the lab.

6) Specimens should be shipped on Monday – Thursday only. Notices will be sent via email

to coordinators one month prior to holidays to indicate any changes to lab and shipping

schedules.

7) Specimens should be shipped when you accumulate 3 subjects. However, no sites should

hold onto samples for more than 4 days; therefore, shippers may need to be sent with fewer

than 3 samples in order to avoid storing samples for more than 4 days.

8) Once shipments are received in the lab, the received date will be entered on the

COPDGene® website by the lab staff. This will confirm and resolve the shipment.

Problems or delays with the shipments will be handled on a case by case basis with clinical

sites as well as logged on the COPDGene® website.

9) The JHU LAB will return the empty boxes to each site so they can be used for future

shipments.

Note: Specimens may only be shipped Monday through Thursday to ensure arrival at the JHU

lab on a weekday. Samples collected on Friday, Saturday, or Sunday will be shipped during the

following week. Exceptions will be dealt with on a case by case basis.

COPDGene®

Version5.0 June 15, 2009

– 73 –

Manual of Operations

Collection of Blood Samples

12d. Supplies and Ordering

By Whom: JHBR Lab and Clinical Center site technicians

Sample Collection Supplies

Item Name

Purple top plastic

EDTA tubes 10 ml

Red-Gray plastic

SST tubes 10 ml

Butterfly needles

21 G (Preferred)

Butterfly needles

23 G

Butterfly needles

25 G

Tourniquets

(reusable)

Tube adapters

Gauze

Bandage

Alcohol pads

Hamilton Bell™

Vanguard Centrifuge

Tube/Document

Labels

Supplier

VWR

VWR

VWR

VWR

VWR

VWR

VWR

VWR

VWR

VWR

MarketLab

AVERY

Catalog Number

BD366643

BD367985

BD367281

BD367283

BD367285

VT367203

VT367290

82004-740

56612-996

15648-981

JL9576

5267

Cost/Unit

case of 10x100

$270.81

case of 10x100

$543.55

case of 4x50

$249.75

case of 4x50

$249.75

case of 4x50

$249.75

case of 20x25

$186.95

case of 10x100

$317.86

case of 5000

$126.54

case of 12x100

$62.02

case of 15x200

$91.81

$359

Local pricing

To be supplied by

Site

Site

Site

Site

Site

Site

Site

Site

Site

Site

Site

Site

*Note: An alternative centrifuge can be used if available.

12e. Good Laboratory Practices

Purpose: To train all staff members in good laboratory practice and ensure these are followed in

all laboratory activities.

By Whom: All laboratory staff members

•

•

Eating, drinking, smoking, applying cosmetics, and handling contact lenses are prohibited in

the laboratory working areas.

Standard precautions should always be followed. Personal protective equipment (gown,

gloves, eye protection) should be worn in the laboratory when handling and processing

specimens and performing diagnostic testing.

Physical containment devices should be used for all manipulations that may cause splashes or

droplets of infectious materials.

Mouth pipetting is forbidden.

Contaminated materials must be disposed of in appropriate biohazard containers.

•

•

•

COPDGene®

Version5.0 June 15, 2009

– 74 –

Manual of Operations

Collection of Blood Samples

•

Work surfaces must be decontaminated after any spill of potentially dangerous material using

a bleach solution. Work surfaces and equipment should be decontaminated after specimens

are processed.

Personnel must wash their hands often – especially after handling infectious materials, before

leaving the laboratory working areas, and before eating.

Personal protective equipment must be removed before leaving the laboratory.

•

•

COPDGene®

Version5.0 June 15, 2009

– 75 –

Manual of Operations

Collection of Blood Samples

Biorepository Appendices:

IATA Shipping guidelines

Purpose: To ensure that all packages meet IATA guidelines for safe shipping.

By Whom: All staff members involved with shipping samples and reagents.

Disclaimer: This information is provided to each site as a quick reference to IATA shipping

®

guidelines. Not all information provided in this section pertains to COPDGene sites (e.g.,

dry ice/carbon dioxide).

Procedure:

The Clinical Center (or shipper), not Federal Express™ (the transport company), is responsible

for determining the hazard class and properly packaging and marking the hazard information on

the shipment.

DIAGNOSTIC SPECIMENS

A diagnostic specimen is any human or animal material including, but not limited to, excreta,

secreta, blood and its components, tissue fluids, and body parts being transported for research,

diagnosis, investigational activities, disease treatment or prevention.

Diagnostic specimens transported under the IATA regulations are assigned the UN identification

number 3373, and are subject to Packing Instructions 650. Any specimens shipped in dry ice

must be labeled with the UN identification number 1845.

Labeling of shipping box:

•

•

All diagnostic specimens must be labeled with a white diamond on point UN 3373

DIAGNOSTIC SPECIMENS label.

If the shipper contains dry ice, there must be a Class 9 diamond on point label on the

face of the shipper, as well as a DRY ICE UN 1845 label that includes the approximate

weight (in kgs) of the dry ice included.

The consignee and the shipper must be identified clearly on the face of the box.

•

PACKING INSTRUCTION 650

STATE VARIATIONS: DOG-03

OPERATOR VARIATIONS: AF-04, AO-03, AS-08, CO-07, CS-07, FX-09, LA-07, LH-12, OF-03

General Requirements

Diagnostic specimens must be packed in good quality packaging, which must be strong enough

to withstand the shocks and loadings normally encountered during transport, including trans-

shipment between transport units and warehouses as well as any removal from a pallet or over

pack for subsequent manual or mechanical handling. Packaging must be constructed and

COPDGene®

Version5.0 June 15, 2009

– 76 –

Manual of Operations

Biorepository Appendix

closed so as to prevent any loss of contents when prepared for transport that might be caused

under normal conditions of transport, by vibration, or by changes in temperature, humidity or

pressure.

Primary receptacles must be packed in secondary packaging in such a way that, under normal

conditions of transport, they cannot break, be punctured or leak their contents into the

secondary packaging. Secondary packaging must be secured in outer packaging with suitable

cushioning material. Any leakage of the contents must not substantially impair the protective

properties of the cushioning material or of the outer packaging. Packages must be prepared as

follows:

(a) For Liquids:

• The primary receptacle(s) must be leak proof and must not contain more than 500 ml;

• There must be absorbent material placed between the primary receptacle and the

secondary packaging; if several fragile primary receptacles are placed in a single

secondary packaging, they must be either individually wrapped or separated so as to

prevent contact between them. The absorbent material, such as cotton wool, must be in

sufficient quantity to absorb the entire contents of the primary receptacles and there

must be a secondary packaging that must be leak proof.

• The primary receptacle or the secondary packaging must be capable of withstanding,

without leakage, an internal pressure producing a pressure differential of not less than

95 kPa in the range of -40°C to +55°C (-40°F to 130°F).

• The outer packaging must not contain more than 4 L.

(b) For Solids:

• The primary receptacle(s) must be sift-proof and must not contain more than 500 g.

• If several fragile primary receptacles are placed in a single secondary packaging, they

must be either individually wrapped or separated so as to prevent contact between them

and there must be a secondary packaging that must be leak proof.

• The outer packaging must not contain more than 4 kg.

An itemized list of contents must be enclosed between the secondary packaging and the outer

packaging.

Each completed package must be capable of successfully passing the drop test described in

6.6.1 except that the height of the drop must not be less than 1.2 m.

Packages must have one side with dimensions of not less than 100 mm x 100 mm (4 in x 4 in)

or packages must be in an over pack that has one side with dimensions of not less than 100

mm x 100 mm (4 in x 4 in).

Each package and the “Nature and Quantity of Goods” box of the air waybill must show the text

"DIAGNOSTIC SPECIMENS." Each package may also be marked in accordance with 7.1.5.8 to

indicate that the shipper has determined that the packaging meets the applicable air transport

requirements. The marking must be applied adjacent to the words "Diagnostic Specimens."

A Shipper’s Declaration for Dangerous Goods is not required.

Provided diagnostic specimens are packed in accordance with this Packing Instruction, no other

requirements of these Regulations apply except for the definition in 3.6.2.1.4 and the reporting

COPDGene®

Version5.0 June 15, 2009

– 77 –

Manual of Operations

Biorepository Appendix

of dangerous goods accidents and incidents in 9.6.1. However, where carbon dioxide, solid (dry

ice) or liquid nitrogen is used to keep specimens cold, all applicable requirements of these

Regulations must be met.

Substances shipped refrigerated or frozen (wet ice, prefrozen packs, Carbon dioxide, solid [dry

ice]): Ice Carbon dioxide, solid (dry ice) or other refrigerant must be placed outside the

secondary packaging(s) or alternatively in an over pack with one or more completed packages.

Interior support must be provided to secure the secondary packaging(s) or packages in the

original position after the ice or Carbon dioxide, solid (dry ice) has been dissipated. If ice is used

the packaging must be leak-proof. If Carbon dioxide, solid (dry ice) is used the outer packaging

must permit the release of carbon-dioxide gas. The primary receptacle must maintain its

containment integrity at the temperature of the refrigerant as well as at the temperatures and

pressure of air transport to which the receptacle could be subjected if refrigeration were to be

lost.

Substances shipped in liquid nitrogen: Plastic capable of withstanding very low temperatures

must be used instead of glass receptacles. Secondary packaging must also withstand very low

temperatures and in most cases will need to be fitted over individual primary receptacles. If

multiple primary receptacles are placed in a single secondary packaging, they must be

separated and supported to ensure that contact between them is prevented. Requirements for

shipment of liquid nitrogen must also be observed. The primary receptacle must maintain its

containment integrity at the temperature of the refrigerant 'used as well as at the temperatures

and pressure of air transport to which the receptacle could be subjected if refrigeration were to

be lost.

Dry ice/carbon dioxide, solid

PACKING INSTRUCTION 904.

STATE VARIATIONS: USG-13

OPERATOR VARIATIONS: HP-02, IC-08, VN-11

Carbon dioxide, solid (dry ice), when offered for transport by air, must be in packaging designed

and constructed to permit the release of carbon dioxide gas and to prevent a build-up of

pressure that could rupture the packaging.

The net weight of the Carbon dioxide, solid (dry ice) must be marked on the outside of the

package.

COPDGene®

Version5.0 June 15, 2009

– 78 –

Manual of Operations

Biorepository Appendix

Step by Step Diagrammatic representation of packaging and shipping

COPDGene® samples using STP-309 shipper

STEP 1.

• Items needed for blood shipment are (clockwise from upper left hand corner):

– Styrofoam insulator box insert and top

– COPDGene barcode labels

– Saf-T-Pak Shipping Box STP-309

– Absorbent Material

– Blood Transmittal Form

– 2 FROZEN Cold Packs

– 3 EDTA and 1 SST 10 ml plastic blood tubes with COPDGene™ barcode

labels

– Bubble Wrap Sleeves for Blood Tubes

– Tyvek Bag

– Plastic Biohazard Bag

COPDGene®

Version5.0 June 15, 2009

– 79 –

Manual of Operations

Biorepository Appendix

STEP 2.

STEP 3.

Before Shipping Blood

REMEMBER:

• All blood tubes must

have the correct

COPDGene™

barcode label

• The SST tube must

be centrifuged

before shipping

• Cold Packs must be

frozen solid in a -20°

centigrade freezer for

at least eight hours prior

to use

• Cold packs should be

placed FLAT in the

freezer for optimal

packing

STEP 4.

Prepare the Shipping

Container:

1. Insert Styrofoam insulator

into STP-309 box

2. Put one cold pack at the

bottom of the Styrofoam

STEP 5.

Insert each blood tube

into an individual

sleeve in the bubble

wrap

STEP 6.

STEP 7.

1

Empty

2

3

4

5

After placing tubes into

the bubble wrap

sleeves, roll up the

bubble wrap and

tubes.

Notice that one bubble wrap sleeve can

hold 5 tubes. Each bubble wrap should

hold one subject’s blood tubes—there will

be one empty sleeve.

COPDGene®

Version5.0 June 15, 2009

– 80 –

Manual of Operations

Biorepository Appendix

STEP 8.

STEP 9.

Secure the blood tubes

in the bubble wrap

with a rubber band

Insert the blood tubes in

the bubble wrap into

the Biohazard bag.

Also insert a loose piece

of the absorbent

material into the

Biohazard bag. It

does not need to the

wrapped around the

blood tubes.

STEP 10.

STEP11.

Putting Three Sets of Blood into

One Biohazard Bag Insert Tubes Here

• When you put three

sets of blood into

one shipment,

space is very tight.

• To ensure that all

materials fit, we

suggest that you

pack the biohazard

bag as shown.

Biohazard Bag

• Ensure that all air is removed from the biohazard bag

before sealing

• Seal the biohazard bag

STEP 12.

STEP 13.

Insert the Biohazard

bag into the Tyvek

bag

• Ensure that all air is

removed from the

Tyvek bag

• Seal the Tyvek Bag

COPDGene®

Version5.0 June 15, 2009

– 81 –

Manual of Operations

Biorepository Appendix

STEP 14.

STEP 15.

• Put the sealed

Tyvek bag on top of

the first frozen cold

pack inside the

shipping container

Place the second

frozen cold pack on

top of the sealed

Tyvek bag

STEP 16.

STEP 17.

Place the Styrofoam

cover on the

container

Place the completed

Blood Transmittal

Form on top of the

Styrofoam container

STEP 18.

STEP 19.

Place the Styrofoam

cover on the

container

Place the completed

Blood Transmittal

Form on top of the

Styrofoam container

COPDGene®

Version5.0 June 15, 2009

– 82 –

Manual of Operations

Biorepository Appendix

STEP 20.

STEP 21.

Place the Styrofoam

cover on the

container

Place the completed

Blood Transmittal

Form on top of the

Styrofoam container

STEP 22.

STEP 23.

Seal the box with

packaging tape

Be sure that the box is compliant with label

regulations.

All boxes are required to have the Biological

Substances B, UN3373 label.

STEP 24.

STEP 25.

• Secure pre-filled FedEx label to the top of

the box.

• Deposit at FedEx drop-off location

• Record shipment immediately on COPDGene website:

https://biosweb.njc.org/sec/COPDGene/sm/trackShipm

ents.cfm?TrackType=blood

COPDGene®

Version5.0 June 15, 2009

– 83 –

Manual of Operations

Biorepository Appendix

Holiday Calendar

2008

New Year's Day

Tuesday, January 1

Do not ship on Dec 24 – Jan 2

Monday, January 21

Monday, February 18

Monday, May 26

Friday, July 4

Do not ship on July 3

Monday, September 1

Thursday, November 27

Do not ship on November 26

Friday, November 28

1/2 day during December

Do not ship Dec 22 – Jan 2

Wednesday, December 24

Thursday, December 25

Wednesday, December 31

Martin Luther King's Birthday

President's Day***

Memorial Day

Independence Day

Labor Day

Thanksgiving Day

Day after Thanksgiving

Holiday Preparation

1/2 Day, Christmas Eve

Christmas Day

1/2 Day, New Year's Eve

2009

New Year's Day

Thursday, January 1

COPDGene®

Version5.0 June 15, 2009

– 84 –

Manual of Operations

Biorepository Appendix